Print Format

1) Intake

Sample Tubes

- Clean tubes - spray wipe with ethanol/alcohol and wipe tubes

- Intake with App (can also do during Amp heating)

2) Make 1X Inactivation

Saline Solution

- Add 100X Inactivation Solution to Saline

| Saline 5 mL | Saline 15 mL | Saline 50 mL |
| --- | --- | --- |
| 100X IS 50 μL | 100X IS 150 μL | 100X IS 500 μL |

3) Add 1X Inactivation

Saline Solution to Sample Tubes

- Add 1mL of inactivation saline solution (1XISS) to each sample tube

- One sample tube open at a time

4) Heat to Inactivate Samples

- Water Bath (set at 99°C)

Heat 8 mintues

Cool 10 minutes

- Dry Heat Block for 1.5mL sample tubes (set at 95°C)

Heat 5 mintues

Cool 5 minutes

5) Make

Reaction Mix

- Add CLAMP MM and Primer Soln (PGS) to 1.5mL tube

- Aliquot into Strip8 tubes or 96 well PCR Plate

| 8 Rxns | 16 Rxns | 24 Rxns | 48 Rxns |
| --- | --- | --- | --- |
| PGS 92 μL | PGS 183 μL | PGS 275 μL | PGS 550 μL |
| CLAMP MM 109 μL | CLAMP MM 218 μL | CLAMP MM 328 μL | CLAMP MM 655 μL |

6) Add Sample

to Reaction

- Add 2μL from each sample tube to corresponding well

- Pipet up and down 5 times, blowout in liquid, tip touch

7) Heat Amp Reactions

- Heat at 65ºC for 25 min

- Set personal timer (on phone) for 24 min

8) Log and

Report Results

- Let samples cool for at least 1 minute before taking photo

- Crop and apply vivid filter

- Log Amp Run Form with QR code link