## Controls
WORKING CONTROLS
Control Name | Control Label | Composition | Spike Type | Spike Concentration | Purpose | Expected Result SARS-CoV-2 Primers | Expected Result RNAseP Primers
No Template Control | PI | 1X PBS + Inactiv Soln | None | 0 | Monitor for contamination of all steps | NEG | NEG
Positive Template Control | PI* | 1X PBS + Inactiv Soln + Spike | Twist RNA Positive Control#1 | 10 cp/ul | Monitor for RNAse contamination during purification and functioning of purification and amplification reagents | POS | NEG
Neg Sample Control | U | Neg Sample Matrix diluted 1:10 in 1XPBS + Inactiv Soln | None | 0 | Monitor for presence of sample material | NEG | POS
Pos Sample Control | K* | Neg Sample Matrix diluted 1:10 in 1XPBS + Inactiv Soln + Spike | Twist RNA Positive Control#1 | 10 cp/ul | Monitor for functioning of inactivation, purification, and amplification reagents | POS | POS
POSITIVE CONTROL SOURCES
Name | Vendor | Catalog Number | Concentration | Spike Protocol
Zeptometrix Viron Control | Zeptometrix | NATSARS(COV2)-ST | 1e6/ml | Spiked into raw sample prior to inactivation solution
Twist RNA Control#1 | Twist | 102019 | 1e9/ml | Spiked into inactivated sample prior to purification
Gamma Irridiated Cell Lysate | BEI / ATCC | NR-52287 | 1.75e9/ml | Spiked into raw sample prior to inactivation solution
Heat Inactivated Cell Lysate | BEI / ATCC | NR-52286 | 3.75e8/ml | Spiked into raw sample prior to inactivation solution

## 1A LOD-Preliminary
1A LOD - Preliminary
|  | Sample Type | Anterior Nares Swab
|  | Swab Type | Cleansky |  | Sample Type | Saliva
|  | Primers | AN |  | Primers | AN
|  | Spike Type | Zeptometrix |  | Spike Type | Zeptometrix
| Concentration | Replicate | Result | Concentration | Replicate | Result
| 5 copies/uL | 1 | POSITIVE | 4 copies/uL | 1 | NEGATIVE
|  | 2 | POSITIVE |  | 2 | POSITIVE
|  | 3 | POSITIVE |  | 3 | POSITIVE
| 2.5 copies/uL | 1 | POSITIVE | 2 copies/uL | 1 | POSITIVE
|  | 2 | POSITIVE |  | 2 | POSITIVE
|  | 3 | POSITIVE | 1 copy/uL | 1 | POSITIVE
| 0 copies/ul | 1 | NEGATIVE |  | 2 | POSITIVE
|  |  |  | 0.5 copies/uL | 1 | POSITIVE
[5 empty] | 2 | POSITIVE

## 1B LOD-Confirmatory
1B LOD - Confirmatory
|  | Sample Type | Anterior Nares Swab
|  | Swab Type | Cleansky |  | Sample Type | Saliva
|  | Primers | AN |  | Primers | AN
|  | Spike Type | Zeptometrix |  | Spike Type | Zeptometrix
| Concentration | Replicate | Result | Concentration | Replicate | Result
| 4 copies/uL | 1 | POSITIVE | 1 copy/uL | 1 | POSITIVE
|  | 2 | POSITIVE |  | 2 | POSITIVE
|  | 3 | POSITIVE |  | 3 | NEGATIVE
|  | 4 | POSITIVE |  | 4 | NEGATIVE
|  | 5 | POSITIVE |  | 5 | POSITIVE
|  | 6 | POSITIVE |  | 6 | POSITIVE
|  | 7 | POSITIVE |  | 7 | POSITIVE
|  | 8 | POSITIVE |  | 8 | POSITIVE
|  | 9 | POSITIVE |  | 9 | INVALID
|  | 10 | POSITIVE |  | 10 | INVALID
|  | 11 | POSITIVE |  | 11 | INVALID
|  | 12 | POSITIVE |  | 12 | NEGATIVE
|  | 13 | POSITIVE |  | 13 | POSITIVE
|  | 14 | POSITIVE |  | 14 | NEGATIVE
|  | 15 | POSITIVE |  | 15 | POSITIVE
|  | 16 | POSITIVE |  | 16 | NEGATIVE
|  | 17 | POSITIVE |  | 17 | NEGATIVE
|  | 18 | POSITIVE |  | 18 | POSITIVE
|  | 19 | POSITIVE |  | 19 | POSITIVE
|  | 20 | POSITIVE |  | 20 | POSITIVE
| 2 copies/uL | 1 | POSITIVE | 0 copies/uL | 1 | NEGATIVE
|  | 2 | POSITIVE |  | 2 | NEGATIVE
|  | 3 | POSITIVE |  | 3 | NEGATIVE
|  | 4 | POSITIVE |  | 4 | NEGATIVE
|  | 5 | POSITIVE |  | 5 | NEGATIVE
|  | 6 | POSITIVE |  | 6 | NEGATIVE
|  | 7 | POSITIVE |  | 7 | NEGATIVE
|  | 8 | POSITIVE |  | 8 | NEGATIVE
|  | 9 | POSITIVE
|  | 10 | POSITIVE | Summary
|  | 11 | POSITIVE | Concentration | Replicates | Result
|  | 12 | POSITIVE | 1 copy/uL | 11 / 20 | POSITIVE
|  | 13 | POSITIVE |  | 6 / 20 | NEGATIVE
|  | 14 | POSITIVE |  | 3 / 20 | INVALID
|  | 15 | POSITIVE | 0 copies/uL | 8 / 8 | NEGATIVE
|  | 16 | POSITIVE
|  | 17 | POSITIVE
|  | 18 | POSITIVE
|  | 19 | POSITIVE
| 0 copies/uL | 1 | NEGATIVE
|  | 2 | NEGATIVE
|  | 3 | NEGATIVE
|  | 4 | NEGATIVE
Summary
| Concentration | Replicates | Result
| 4 copies/uL | 20 / 20 | POSITIVE
| 2 copies/uL | 19 / 19 | POSITIVE
| 0 copies/uL | 8 / 8 | NEGATIVE

## 2 Inclusivity-In Silico
2 Analytical Reactivity / Inclusivity - In Silico Analysis

## 3 Analytical Specificity
3 Analytical Specificity - Cross Reactivity

## 4 Interfering Substances
4 Analytical Specificity - Interfering Substances

## 5 Clinical Evaluation
5 Clinical Evaluation