## 1A LoD-Preliminary
1A LoD - Preliminary
Plate Map
|  | 1 | 2 | 3
| A Row | 100000 | 100000 | 100000
| B Row | 50000 | 50000 | 50000
| C Row | 25000 | 25000 | 25000
| D Row | 12500 | 12500 | 12500
| E Row | 6250 | 6250 | 6250
| F Row | 3125 | 3125 | 3125
| G Row | 0 | 0 | 0
| H Row | Neg Control | Neg Control | Pos Control
| Line Item Data - Applied Biosystems QuantStudio 7 Pro |  |  |  |  | Summary Data - Applied Biosystems QuantStudio 7 Pro
| Concentration (copies/mL) | Replicate | N1 Ct | RP Ct | Result | Concentration (copies/mL) | N1 Mean Ct | N1 SD Ct | RP Mean Ct | RP SD Ct | N1 Detected
| 100000 | 1 | 31.8 | 27.5 | POSITIVE | 100000 | 31.6667 | 0.11547 | 27.4 | 0.17321 | 100% (3/3)
|  | 2 | 31.6 | 27.2 | POSITIVE | 50000 | 32.8333 | 0.25166 | 27.8333 | 0.20817 | 100% (3/3)
|  | 3 | 31.6 | 27.5 | POSITIVE | 25000 | 33.9667 | 0.11547 | 27.7333 | 0.25166 | 100% (3/3)
| 50000 | 1 | 32.6 | 28 | POSITIVE | 12500 | 34.7 | 0.2 | 27.8 | 0.3 | 100% (3/3)
|  | 2 | 33.1 | 27.9 | POSITIVE | 6250 | 36 | 0.36056 | 28 | 0.1 | 100% (3/3)
|  | 3 | 32.8 | 27.6 | POSITIVE | 3125 | 37.3333 | 0.50332 | 27.9333 | 0.057735 | 100% (3/3)
| 25000 | 1 | 33.9 | 28 | POSITIVE | 0 | Undeterm | N/A | 27.8 | 0.17321 | 0% (3/3)
|  | 2 | 34.1 | 27.5 | POSITIVE
|  | 3 | 33.9 | 27.7 | POSITIVE
| 12500 | 1 | 34.7 | 28.1 | POSITIVE
|  | 2 | 34.5 | 27.5 | POSITIVE
|  | 3 | 34.9 | 27.8 | POSITIVE
| 6250 | 1 | 35.6 | 27.9 | POSITIVE
|  | 2 | 36.1 | 28 | POSITIVE
|  | 3 | 36.3 | 28.1 | POSITIVE
| 3125 | 1 | 37.4 | 28 | POSITIVE
|  | 2 | 37.8 | 27.9 | POSITIVE
|  | 3 | 36.8 | 27.9 | POSITIVE
| 0 | 1 | Undetermined | 27.9 | NEGATIVE
|  | 2 | Undetermined | 27.9 | NEGATIVE
|  | 3 | Undetermined | 27.6 | NEGATIVE
| Negative Control | 1 | Undetermined | Undetermined | NEGATIVE
|  | 2 | Undetermined | Undetermined | NEGATIVE
| Positive Control | 1 | 31.1 | 31 | POSITIVE

## 1B LoD-Confirmatory
1A LoD - Confirmatory
Plate Map
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
| A Row | 1560 | 1560 | 1560 | 3125 | 3125 | 3125 | 6250 | 6250 | 6250
| B Row | 1560 | 1560 | 1560 | 3125 | 3125 | 3125 | 6250 | 6250 | 6250
| C Row | 1560 | 1560 | 1560 | 3125 | 3125 | 3125 | 6250 | 6250 | 6250
| D Row | 1560 | 1560 | 1560 | 3125 | 3125 | 3125 | 6250 | 6250 | 6250
| E Row | 1560 | 1560 | 1560 | 3125 | 3125 | 3125 | 6250 | 6250 | 6250
| F Row | 1560 | 1560 | 1560 | 3125 | 3125 | 3125 | 6250 | 6250 | 6250
| G Row | 1560 | 1560 | 1560 | 3125 | 3125 | 3125 | 6250 | 6250 | 6250
| H Row | 0 | 0 | 0 | Neg Control | Neg Control | Neg Control | Neg Control | Neg Control | Pos Control
| Summary Data - Applied Biosystems QuantStudio 7 Pro | 7: Summary Data - Bio-Rad CFX96 |  |  |  | Summary Data - Applied Biosystems QuantStudio 6 Flex
| Concentration (copies/mL) | N1 Mean Ct | N1 SD Ct | RP Mean Ct | RP SD Ct | Detected < Ct 38 |  | Concentration (copies/mL) | N1 Mean Ct | N1 SD Ct | RP Mean Ct | RP SD Ct | Detected < Ct 40 | Concentration (copies/mL) | N1 Mean Ct | N1 SD Ct | RP Mean Ct | RP SD Ct | Detected < Ct 40
| 6250 | 35.3744 | 0.59012 | 28.8092 | 0.09987 | 100% (21/21) |  | 6250 | 36.0571 | 0.65694 | 27.3286 | 0.23483 | 100% (21/21) | 6250 | 35.6476 | 0.85301 | 26.9524 | 0.12891 | 100% (21/21)
| 3125 | 36.5 | 0.6 | 28.8429 | 0.3187 | 95% (20/21) |  | 3125 | 37.18 | 0.88591 | 27.2571 | 0.10757 | 95% (20/21) | 3125 | 36.9 | 1.0986 | 27.0095 | 0.13002 | 100% (21/21)
| 1560 | 37 | 0.8 | 28.8 | 0.2 | 57% (12/21) |  | 1560 | 38.2368 | 0.76609 | 27.2857 | 0.1526 | 90% (19/21) | 1560 | 37.6056 | 0.73522 | 26.9905 | 0.083095 | 86% (18/21)
| Line Item Data - Applied Biosystems QuantStudio 7 Pro | 7: Line Item Data - Bio-Rad CFX96 |  |  |  | Line Item Data - Applied Biosystems QuantStudio 6 Flex
| Concentration (copies/mL) | Replicate | N1 Ct | RP Ct | Result |  |  | Concentration (copies/mL) | Replicate | N1 Ct | RP Ct | Result |  | Concentration (copies/mL) | Replicate | N1 Ct | RP Ct | Result
| 6250 | 1 | 35.5 | 29.1 | POSITIVE | N1 Mean Ct |  | 6250 | 1 | 35.1 | 28.2 | POSITIVE | N1 Mean Ct | 6250 | 1 | 35.9 | 27 | POSITIVE | N1 Mean Ct
|  | 2 | 35.3 | 28.7 | POSITIVE | 35.3744 |  |  | 2 | 35.5 | 27.3 | POSITIVE | 36.0571 |  | 2 | 35.4 | 26.9 | POSITIVE | 35.6476
|  | 3 | 35.7 | 28.7 | POSITIVE | N1 SD Ct |  |  | 3 | 35.7 | 27.2 | POSITIVE | N1 SD Ct |  | 3 | 35.8 | 27 | POSITIVE | N1 SD Ct
|  | 4 | 35.1 | 28.9 | POSITIVE | 0.59012 |  |  | 4 | 35.9 | 27.4 | POSITIVE | 0.65694 |  | 4 | 35.5 | 26.9 | POSITIVE | 0.85301
|  | 5 | 35.7 | 28.8 | POSITIVE |  |  |  | 5 | 36 | 27.2 | POSITIVE |  |  | 5 | 34.9 | 26.8 | POSITIVE
|  | 6 | 34.6 | 28.8 | POSITIVE | RP Mean Ct |  |  | 6 | 35.6 | 27.2 | POSITIVE | RP Mean Ct |  | 6 | 36.2 | 26.9 | POSITIVE | RP Mean Ct
|  | 7 | 34.9 | 28.9 | POSITIVE | 28.8092 |  |  | 7 | 36 | 27.1 | POSITIVE | 27.3286 |  | 7 | 35.8 | 27 | POSITIVE | 26.9524
|  | 8 | 35.6 | 28.8 | POSITIVE | RP SD Ct |  |  | 8 | 35.3 | 27.4 | POSITIVE | RP SD Ct |  | 8 | 35.1 | 27 | POSITIVE | RP SD Ct
|  | 9 | 35.2 | 28.9 | POSITIVE | 0.09987 |  |  | 9 | 36.2 | 27.2 | POSITIVE | 0.23483 |  | 9 | 36.1 | 27 | POSITIVE | 0.12891
|  | 10 | 35.5 | 28.9 | POSITIVE |  |  |  | 10 | 35.6 | 27.3 | POSITIVE |  |  | 10 | 34.7 | 26.9 | POSITIVE
|  | 11 | 35.8 | 28.7 | POSITIVE | POSITIVE |  |  | 11 | 36.5 | 27.3 | POSITIVE | POSITIVE |  | 11 | 35.1 | 26.8 | POSITIVE | POSITIVE
|  | 12 | 35.5 | 28.8 | POSITIVE | 21 |  |  | 12 | 36.6 | 27.3 | POSITIVE | 21 |  | 12 | 34.9 | 27 | POSITIVE | 21
|  | 13 | 35.5 | 28.8 | POSITIVE |  |  |  | 13 | 37.1 | 27.4 | POSITIVE |  |  | 13 | 35.5 | 27 | POSITIVE
|  | 14 | 34 | 28.8 | POSITIVE | NEGATIVE |  |  | 14 | 37.5 | 27.2 | POSITIVE | NEGATIVE |  | 14 | 34.8 | 26.9 | POSITIVE | NEGATIVE
|  | 15 | 35.6 | 28.8 | POSITIVE | 0 |  |  | 15 | 35.1 | 27.6 | POSITIVE | 0 |  | 15 | 36.6 | 27 | POSITIVE | 0
|  | 16 | 35.1 | 28.8 | POSITIVE |  |  |  | 16 | 36.5 | 27.5 | POSITIVE |  |  | 16 | 35.7 | 27 | POSITIVE
|  | 17 | 35.4 | 28.8 | POSITIVE |  |  |  | 17 | 36.3 | 27.2 | POSITIVE |  |  | 17 | 35 | 26.8 | POSITIVE
|  | 18 | 34.7 | 28.7 | POSITIVE |  |  |  | 18 | 36 | 27.2 | POSITIVE |  |  | 18 | 35.3 | 26.8 | POSITIVE
|  | 19 | 35 | 28.9 | POSITIVE |  |  |  | 19 | 36.7 | 27.3 | POSITIVE |  |  | 19 | 38.6 | 27.4 | POSITIVE
|  | 20 | 36.6 | 28.7 | POSITIVE |  |  |  | 20 | 35.3 | 27.1 | POSITIVE |  |  | 20 | 36.2 | 27 | POSITIVE
|  | 21 | 36.5614 | 28.6938 | POSITIVE |  |  |  | 21 | 36.7 | 27.3 | POSITIVE |  |  | 21 | 35.5 | 26.9 | POSITIVE
| 3125 | 1 | 35.6 | 28.5 | POSITIVE | N1 Mean Ct |  | 3125 | 1 | 36.4 | 27.1 | POSITIVE | N1 Mean Ct | 3125 | 1 | 37.3 | 27.2 | POSITIVE | N1 Mean Ct
|  | 2 | 36.1 | 28.8 | POSITIVE | 36.45 |  |  | 2 | 36.6 | 27.2 | POSITIVE | 37.18 |  | 2 | 38.7 | 27.2 | POSITIVE | 36.9
|  | 3 | 36.5 | 28.8 | POSITIVE | N1 SD Ct |  |  | 3 | 35.7 | 27.3 | POSITIVE | N1 SD Ct |  | 3 | 37.4 | 27 | POSITIVE | N1 SD Ct
|  | 4 | 35.6 | 28.5 | POSITIVE | 0.61516 |  |  | 4 | 37.9 | 27.1 | POSITIVE | 0.88591 |  | 4 | 36.7 | 26.9 | POSITIVE | 1.0986
|  | 5 | 35.7 | 28.7 | POSITIVE |  |  |  | 5 | 36.9 | 27.3 | POSITIVE |  |  | 5 | 36.3 | 27 | POSITIVE
|  | 6 | 38.6 | 29 | NEGATIVE | RP Mean Ct |  |  | 6 | 36.5 | 27.4 | POSITIVE | RP Mean Ct |  | 6 | 35.8 | 27.1 | POSITIVE | RP Mean Ct
|  | 7 | 37 | 28.6 | POSITIVE | 28.835 |  |  | 7 | 37.8 | 27.3 | POSITIVE | 27.26 |  | 7 | 36.1 | 27.1 | POSITIVE | 27.0095
|  | 8 | 35.7 | 28.6 | POSITIVE | RP SD Ct |  |  | 8 | 36.7 | 27.4 | POSITIVE | RP SD Ct |  | 8 | 36.3 | 27.3 | POSITIVE | RP SD Ct
|  | 9 | 36.3 | 28.9 | POSITIVE | 0.32489 |  |  | 9 | 36.3 | 27.2 | POSITIVE | 0.10954 |  | 9 | 37.7 | 27 | POSITIVE | 0.13002
|  | 10 | 36.6 | 28.7 | POSITIVE |  |  |  | 10 | 39 | 27.1 | POSITIVE |  |  | 10 | 36.5 | 27 | POSITIVE
|  | 11 | 37.5 | 28.8 | POSITIVE | POSITIVE |  |  | 11 | 37.9 | 27.3 | POSITIVE | POSITIVE |  | 11 | 36.1 | 27 | POSITIVE | POSITIVE
|  | 12 | 36.9 | 28.9 | POSITIVE | 20 |  |  | 12 | 36.6 | 27.1 | POSITIVE | 20 |  | 12 | 35.5 | 26.9 | POSITIVE | 21
|  | 13 | 37.6 | 28.6 | POSITIVE |  |  |  | 13 | 36.2 | 27.3 | POSITIVE |  |  | 13 | 38.8 | 26.8 | POSITIVE
|  | 14 | 37 | 29.7 | POSITIVE | NEGATIVE |  |  | 14 | 37.7 | 27.3 | POSITIVE | NEGATIVE |  | 14 | 36.3 | 26.8 | POSITIVE | NEGATIVE
|  | 15 | 36.7 | 29.2 | POSITIVE | 1 |  |  | 15 | 36.8 | 27.1 | POSITIVE | 1 |  | 15 | 36.4 | 27.1 | POSITIVE | 0
|  | 16 | 35.6 | 28.7 | POSITIVE |  |  |  | 16 | 37.9 | 27.3 | POSITIVE |  |  | 16 | 36.7 | 27.1 | POSITIVE
|  | 17 | 36.5 | 28.9 | POSITIVE |  |  |  | 17 | 38.7 | 27.3 | POSITIVE |  |  | 17 | 35.9 | 26.9 | POSITIVE
|  | 18 | 37 | 29.6 | POSITIVE |  |  |  | 18 | 38.1 | 27.4 | POSITIVE |  |  | 18 | 37.7 | 27 | POSITIVE
|  | 19 | 36.5 | 28.7 | POSITIVE |  |  |  | 19 | Undetermined | 27.2 | NEGATIVE |  |  | 19 | 35.6 | 26.9 | POSITIVE
|  | 20 | 36.1 | 28.6 | POSITIVE |  |  |  | 20 | 36.9 | 27.3 | POSITIVE |  |  | 20 | 39.4 | 26.9 | POSITIVE
|  | 21 | 36.5 | 28.9 | POSITIVE |  |  |  | 21 | 37 | 27.4 | POSITIVE |  |  | 21 | 37.7 | 27 | POSITIVE
| 1560 | 1 | 37.6 | 28.7 | POSITIVE | N1 Mean Ct |  | 1560 | 1 | Undetermined | 27.1 | NEGATIVE | N1 Mean Ct | 1560 | 1 | 36.5 | 27.1 | POSITIVE | N1 Mean Ct
|  | 2 | Undetermined | 29.3 | NEGATIVE | 36.9917 |  |  | 2 | 37.9 | 27.2 | POSITIVE | 38.2368 |  | 2 | 37.6 | 27.1 | POSITIVE | 37.6056
|  | 3 | 38.5 | 28.8 | NEGATIVE | N1 SD Ct |  |  | 3 | 38.8 | 27.2 | POSITIVE | N1 SD Ct |  | 3 | 37.7 | 27 | POSITIVE | N1 SD Ct
|  | 4 | 37.3 | 28.7 | POSITIVE | 0.76212 |  |  | 4 | 36.8 | 27.1 | POSITIVE | 0.76609 |  | 4 | 37.1 | 27.1 | POSITIVE | 0.73522
|  | 5 | 37.1 | 29 | POSITIVE |  |  |  | 5 | 38.4 | 27.2 | POSITIVE |  |  | 5 | 37.3 | 27 | POSITIVE
|  | 6 | 37.6 | 28.8 | POSITIVE | RP Mean Ct |  |  | 6 | 39 | 27.1 | POSITIVE | RP Mean Ct |  | 6 | Undetermined | 27.1 | NEGATIVE | RP Mean Ct
|  | 7 | 39.1 | 28.7 | NEGATIVE | 28.7583 |  |  | 7 | 37.9 | 27.2 | POSITIVE | 27.2895 |  | 7 | 36.9 | 27.1 | POSITIVE | 26.9833
|  | 8 | 36.2 | 29 | POSITIVE | RP SD Ct |  |  | 8 | Undetermined | 27.4 | NEGATIVE | RP SD Ct |  | 8 | 37 | 27 | POSITIVE | RP SD Ct
|  | 9 | 37.6 | 28.8 | POSITIVE | 0.19287 |  |  | 9 | 38.7 | 27.2 | POSITIVE | 0.15237 |  | 9 | 37.7 | 26.9 | POSITIVE | 0.085749
|  | 10 | 37.1 | 28.5 | POSITIVE |  |  |  | 10 | 38.1 | 27.3 | POSITIVE |  |  | 10 | 38.7 | 26.8 | POSITIVE
|  | 11 | 35.9 | 29.1 | POSITIVE | POSITIVE |  |  | 11 | 39.4 | 27.5 | POSITIVE | POSITIVE |  | 11 | 36.6 | 27 | POSITIVE | POSITIVE
|  | 12 | 38.5 | 28.9 | NEGATIVE | 12 |  |  | 12 | 36.6 | 27.4 | POSITIVE | 19 |  | 12 | 37.6 | 27 | POSITIVE | 18
|  | 13 | 36.3 | 28.5 | POSITIVE |  |  |  | 13 | 38.3 | 27.5 | POSITIVE |  |  | 13 | Undetermined | 27 | NEGATIVE
|  | 14 | 38.4 | 30.1 | NEGATIVE | NEGATIVE |  |  | 14 | 38.2 | 27.5 | POSITIVE | NEGATIVE |  | 14 | 38.6 | 27 | POSITIVE | NEGATIVE
|  | 15 | 37.6 | 28.7 | POSITIVE | 9 |  |  | 15 | 38.3 | 27.6 | POSITIVE | 2 |  | 15 | Undetermined | 27 | NEGATIVE | 3
|  | 16 | 37.9 | 28.7 | POSITIVE |  |  |  | 16 | 39.2 | 27.3 | POSITIVE |  |  | 16 | 38.1 | 26.9 | POSITIVE
|  | 17 | 38.6 | 29 | NEGATIVE |  |  |  | 17 | 39.2 | 27.2 | POSITIVE |  |  | 17 | 36.8 | 26.9 | POSITIVE
|  | 18 | 38.6 | 28.9 | NEGATIVE |  |  |  | 18 | 37.6 | 27.3 | POSITIVE |  |  | 18 | 38.7 | 27 | POSITIVE
|  | 19 | 35.7 | 28.6 | POSITIVE |  |  |  | 19 | 37.6 | 27.4 | POSITIVE |  |  | 19 | 37.5 | 26.9 | POSITIVE
|  | 20 | Undetermined | 29.1 | NEGATIVE |  |  |  | 20 | 38.7 | 27.1 | POSITIVE |  |  | 20 | 38.8 | 26.9 | POSITIVE
|  | 21 | 38.6 | 28.9 | NEGATIVE |  |  |  | 21 | 37.8 | 27.2 | POSITIVE |  |  | 21 | 37.7 | 27 | POSITIVE
| 0 | 1 | Undetermined | 28.6441 | NEGATIVE |  |  | 0 | 1 | Undetermined | 27.4 | NEGATIVE |  | 0 | 1 | Undetermined | 27 | NEGATIVE
|  | 2 | Undetermined | 28.6231 | NEGATIVE |  |  |  | 2 | Undetermined | 27.3 | NEGATIVE |  |  | 2 | Undetermined | 27.4 | NEGATIVE
|  | 3 | Undetermined | 28.459 | NEGATIVE |  |  |  | 3 | 38.1 | 27.5 | POSITIVE |  |  | 3 | Undetermined | 26.9 | NEGATIVE
| Neg Control | 1 | Undetermined | Undetermined | NEGATIVE |  |  | Neg Control | 1 | Undetermined | Undetermined | NEGATIVE |  | Neg Control | 1 | Undetermined | Undetermined | NEGATIVE
|  | 2 | Undetermined | Undetermined | NEGATIVE |  |  |  | 2 | Undetermined | Undetermined | NEGATIVE |  |  | 2 | Undetermined | Undetermined | NEGATIVE
|  | 3 | Undetermined | Undetermined | NEGATIVE |  |  |  | 3 | Undetermined | Undetermined | NEGATIVE |  |  | 3 | Undetermined | Undetermined | NEGATIVE
|  | 4 | Undetermined | Undetermined | NEGATIVE |  |  |  | 4 | Undetermined | 40.5 | NEGATIVE |  |  | 4 | Undetermined | Undetermined | NEGATIVE
|  | 5 | Undetermined | Undetermined | NEGATIVE |  |  |  | 5 | Undetermined | 41.2 | NEGATIVE |  |  | 5 | Undetermined | Undetermined | NEGATIVE
| Pos Control | 1 | 31.018 | 31.3939 | POSITIVE |  |  | Pos Control | 1 | 31.1 | 31.2 | POSITIVE |  | Pos Control | 1 | 30.5 | 31 | POSITIVE

## 3C Interfering Substances
3C Interfering Substances
Plate Map
|  | 1 | 2 | 3 | 4 | 5 | 6
| A Row | INF1* | INF1* | INF1* | INF10* | INF10* | INF10*
| B Row | INF1 | INF1 | INF1 | INF10 | INF10 | INF10
| C Row | INF3* | INF3* | INF3* | Positive Control* | Positive Control* | Positive Control*
| D Row | INF3 | INF3 | INF3 | NOINF* | NOINF* | NOINF*
| E Row | INF4* | INF4* | INF4* | Neg Control | Pos Control | Pos Control
| F Row | 1: INF4; 2: INF4; 3: INF4
| G Row | 1: INF9*; 2: INF9*; 3: INF9*
| H Row | 1: INF9; 2: INF9; 3: INF9
Summary Data
| Specimen ID | Interferring Substance | Active Ingredient | Concentration | Positive Control N1 |  | Negative Control N1 |  | Positive Control RP |  | Negative Control RP |  | % Agreement with Expected Results
|  | 4: Mean Ct; 5: SD Ct; 6: Mean Ct | SD Ct | Mean Ct | SD Ct | Mean Ct | SD Ct | Positive Control Spiked | Negative Control Unspiked
| INF1 | Blood | N/A | 1% v/v | 36.4303 | 0.35055 | Undetermined | n/a | 32.5 | 0.1 | 29.4667 | 0.15275 | 100% (3/3) | 100% (3/3)
| INF2 | Nasal Congestion Spray | Acetaminophen, Guaifenesin, Phenylephrine HCI | 20% v/v | 37.1916 | 0.51912 | Undetermined | n/a | 31.8 | 0.2 | 28.8 | 0.36056 | 100% (3/3) | 100% (3/3)
| INF3 | Nasal Allergy Spray | Oxymetazoline HCl | 15% v/v | 37.2002 | 0.47563 | Undetermined | n/a | 30.1333 | 0.15275 | 30 | 0 | 100% (3/3) | 100% (3/3)
| INF4 | Lozenges | Menthol | 10% w/v | 36.002 | 0.37601 | Undetermined | n/a | 30.8667 | 0.25166 | 27.8667 | 0.11547 | 100% (3/3) | 100% (3/3)
| INF5 | Mucin | N/A | 0.5% w/v | 35.8166 | 0.14632 | Undetermined | n/a | 33.1 | 0.17321 | 31.2667 | 0.11547 | 100% (3/3) | 100% (3/3)
Line Item Data - Applied Biosystems QuantStudio 7 Pro
| Specimen ID | Interferring Substance | Spiked? | Replicate | N1 Ct | RP Ct | Result
| INF1* | Blood | Spiked | 1 | 36.1573 | 32.5 | POSITIVE
|  |  |  | 2 | 36.8256 | 32.6 | POSITIVE
|  |  |  | 3 | 36.308 | 32.4 | POSITIVE
| INF1 | Blood | Unspiked | 1 | Undetermined | 29.6 | NEGATIVE
|  |  |  | 2 | Undetermined | 29.3 | NEGATIVE
|  |  |  | 3 | Undetermined | 29.5 | NEGATIVE
| INF2* | Nasal Congestion Spray | Spiked | 1 | 37.4919 | 31.6 | POSITIVE
|  |  |  | 2 | 36.5922 | 31.8 | POSITIVE
|  |  |  | 3 | 37.4907 | 32 | POSITIVE
| INF2 | Nasal Congestion Spray | Unspiked | 1 | Undetermined | 28.5 | NEGATIVE
|  |  |  | 2 | Undetermined | 29.2 | NEGATIVE
|  |  |  | 3 | Undetermined | 28.7 | NEGATIVE
| INF3* | Nasal Allergy Spray | Spiked | 1 | 37.307 | 30.1 | POSITIVE
|  |  |  | 2 | 37.6134 | 30 | POSITIVE
|  |  |  | 3 | 36.6802 | 30.3 | POSITIVE
| INF3 | Nasal Allergy Spray | Unspiked | 1 | Undetermined | 30 | NEGATIVE
|  |  |  | 2 | Undetermined | 30 | NEGATIVE
|  |  |  | 3 | Undetermined | 30 | NEGATIVE
| INF4* | Lozenges | Spiked | 1 | 35.7641 | 30.6 | POSITIVE
|  |  |  | 2 | 36.4355 | 31.1 | POSITIVE
|  |  |  | 3 | 35.8063 | 30.9 | POSITIVE
| INF4 | Lozenges | Unspiked | 1 | Undetermined | 28 | NEGATIVE
|  |  |  | 2 | Undetermined | 27.8 | NEGATIVE
|  |  |  | 3 | Undetermined | 27.8 | NEGATIVE
| INF5* | Mucin | Spiked | 1 | 35.9646 | 33 | POSITIVE
|  |  |  | 2 | 35.6721 | 33 | POSITIVE
|  |  |  | 3 | 35.813 | 33.3 | POSITIVE
| INF5 | Mucin | Unspiked | 1 | Undetermined | 31.2 | NEGATIVE
|  |  |  | 2 | Undetermined | 31.2 | NEGATIVE
|  |  |  | 3 | Undetermined | 31.4 | NEGATIVE
| INFNONE* | None | Spiked | 1 | 36.6836 | 28.6 | POSITIVE
|  |  |  | 2 | 37.4144 | 28.9 | POSITIVE
|  |  |  | 3 | 37.2554 | 28.8 | POSITIVE
| Positive Control* | None | Spiked | 1 | 36.4141 | Undetermined | POSITIVE
|  |  |  | 2 | 35.3738 | Undetermined | POSITIVE
|  |  |  | 3 | 35.8054 | Undetermined | POSITIVE
| Positive Control | None | Unspiked | 1 | 37.4726 | 27.7 | POSITIVE
|  |  |  | 2 | 36.9066 | 27.7 | POSITIVE
| Negative Control | None | Unspiked | 1 | Undetermined | Undetermined | NEGATIVE

## 4 Clinical Evaluation
4 Clinical Evaluation
Plate Map
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
| A Row | ST01 | ST09 | ST17 | ST25 | ST33 | ST41 | ST49 | ST57 | ST65 | ST73 | Pos Control
| B Row | ST02 | ST10 | ST18 | ST26 | ST34 | ST42 | ST50 | ST58 | ST66 | ST74 | Neg Control
| C Row | ST03 | ST11 | ST19 | ST27 | ST35 | ST43 | ST51 | ST59 | ST67 | ST75
| D Row | ST04 | ST12 | ST20 | ST28 | ST36 | ST44 | ST52 | ST60 | ST68 | ST76
| E Row | ST05 | ST13 | ST21 | ST29 | ST37 | ST45 | ST53 | ST61 | ST69 | ST77
| F Row | ST06 | ST14 | ST22 | ST30 | ST38 | ST46 | ST54 | ST62 | ST70 | ST78
| G Row | ST07 | ST15 | ST23 | ST31 | ST39 | ST47 | ST55 | ST63 | ST71 | ST79
| H Row | ST08 | ST16 | ST24 | ST32 | ST40 | ST48 | ST56 | ST64 | ST72 | ST80
Summary Data
| FloodLAMP SwabDirect PCR Result | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: Comparator Positive; 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: Comparator Negative; 3 ST17 ST18 ST19 ST20 ST21 ST22 ST23 ST24: Total
| Positive | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 39; 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: 0; 3 ST17 ST18 ST19 ST20 ST21 ST22 ST23 ST24: 39
| Negative | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 1; 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: 40; 3 ST17 ST18 ST19 ST20 ST21 ST22 ST23 ST24: 41
| Invalid | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 0; 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: 0; 3 ST17 ST18 ST19 ST20 ST21 ST22 ST23 ST24: 0
| Total | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 40; 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: 40; 3 ST17 ST18 ST19 ST20 ST21 ST22 ST23 ST24: 80
| Positive Agreement | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 97.5% (39/40) 95% CI: 86.8% to 99.9%
| Negative Agreement | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 100% (40/40) 95% CI: 91.2% to 100%
Source of Specimens: Stanford COVID-19 Clinical Testing Program
Specimen Type: Anterior Nares Swab in PBS, previously tested and frozen
Comparator Test: Hologic Panther Fusion SARS-CoV-2 Assay and Hologic Panther Aptima SARS-CoV-2 Assay
| Line Item Data | 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: Comparator - Hologic Panther; 6 ST41 ST42 ST43 ST44 ST45 ST46 ST47 ST48: FloodLAMP SwabDirect PCR - Bio-Rad CFX96
| Well | Specimen ID | Comparator Test Date | Comparator Test Assay | Comparator Ct | Comparator Result | N1 Ct | RP Ct | Result
| D01 | ST04 | 2021-01-02 | Panther Fusion | 13.2 | POSITIVE | 22.14 | Und | POSITIVE
| C05 | ST35 | 2021-01-02 | Panther Fusion | 13.6 | POSITIVE | 23.36 | Und | POSITIVE
| H07 | ST56 | 2021-01-02 | Panther Fusion | 14 | POSITIVE | 23.72 | Und | POSITIVE
| G03 | ST23 | 2021-01-01 | Panther Fusion | 15 | POSITIVE | 20.48 | Und | POSITIVE
| C04 | ST27 | 2021-01-04 | Panther Fusion | 15.1 | POSITIVE | 21.42 | Und | POSITIVE
| E09 | ST69 | 2021-01-03 | Panther Fusion | 15.4 | POSITIVE | 23.94 | 44.84 | POSITIVE
| F02 | ST14 | 2021-01-03 | Panther Fusion | 15.5 | POSITIVE | 22.92 | Und | POSITIVE
| B02 | ST10 | 2021-01-01 | Panther Fusion | 15.6 | POSITIVE | 27.68 | 29.88 | POSITIVE
| E07 | ST53 | 2021-01-02 | Panther Fusion | 15.6 | POSITIVE | 25.09 | 26.37 | POSITIVE
| D08 | ST60 | 2021-01-03 | Panther Fusion | 15.7 | POSITIVE | 21.65 | 39.12 | POSITIVE
| A04 | ST25 | 2021-01-01 | Panther Fusion | 16.5 | POSITIVE | 26.53 | 29.28 | POSITIVE
| F03 | ST22 | 2021-01-03 | Panther Fusion | 17.2 | POSITIVE | 24.73 | Und | POSITIVE
| G02 | ST15 | 2021-01-01 | Panther Fusion | 17.4 | POSITIVE | 29.13 | 28.91 | POSITIVE
| B03 | ST18 | 2021-01-02 | Panther Fusion | 17.5 | POSITIVE | 27.71 | 28.42 | POSITIVE
| H03 | ST24 | 2021-01-01 | Panther Fusion | 18.9 | POSITIVE | 30.06 | 29.81 | POSITIVE
| E06 | ST45 | 2021-01-06 | Panther Fusion | 18.9 | POSITIVE | 25.17 | 34.74 | POSITIVE
| G08 | ST63 | 2021-01-02 | Panther Fusion | 19.4 | POSITIVE | 26.19 | 27.69 | POSITIVE
| D07 | ST52 | 2021-01-01 | Panther Fusion | 19.8 | POSITIVE | 29.97 | 28.32 | POSITIVE
| H06 | ST48 | 2021-01-02 | Panther Fusion | 20 | POSITIVE | 32.66 | 32.43 | POSITIVE
| E04 | ST29 | 2021-01-01 | Panther Fusion | 20.5 | POSITIVE | 30.69 | 29.67 | POSITIVE
| B09 | ST66 | 2021-01-01 | Panther Fusion | 21.1 | POSITIVE | 28.81 | 28.38 | POSITIVE
| F01 | ST06 | 2021-01-05 | Panther Fusion | 22.6 | POSITIVE | 29.72 | 27.85 | POSITIVE
| F05 | ST38 | 2021-01-01 | Panther Fusion | 23.2 | POSITIVE | 32.03 | 27.71 | POSITIVE
| F04 | ST30 | 2021-01-09 | Panther Fusion | 23.7 | POSITIVE | 31.05 | 31.98 | POSITIVE
| D10 | ST76 | 2021-01-02 | Panther Fusion | 24.1 | POSITIVE | 32.4 | 28.55 | POSITIVE
| G07 | ST55 | 2021-01-01 | Panther Fusion | 24.8 | POSITIVE | 32.38 | 28.19 | POSITIVE
| D03 | ST20 | 2021-01-01 | Panther Fusion | 25.1 | POSITIVE | 35.05 | 26.4 | POSITIVE
| H10 | ST80 | 2021-01-01 | Panther Fusion | 25.2 | POSITIVE | 36.62 | 30.51 | POSITIVE
| H08 | ST64 | 2021-01-02 | Panther Fusion | 25.6 | POSITIVE | 37.94 | 32.77 | POSITIVE
| H09 | ST72 | 2021-01-02 | Panther Fusion | 25.6 | POSITIVE | 37.26 | 31.28 | POSITIVE
| B10 | ST74 | 2021-01-02 | Panther Fusion | 25.7 | POSITIVE | 35.56 | 31.43 | POSITIVE
| G06 | ST47 | 2021-01-02 | Panther Fusion | 25.9 | POSITIVE | 36.48 | 30.08 | POSITIVE
| A09 | ST65 | 2021-01-05 | Panther Fusion | 26.1 | POSITIVE | 33.86 | 30.38 | POSITIVE
| D02 | ST12 | 2021-01-02 | Panther Fusion | 26.3 | POSITIVE | 36.12 | 32.04 | POSITIVE
| B05 | ST34 | 2021-01-04 | Panther Fusion | 26.4 | POSITIVE | 33.78 | 28.91 | POSITIVE
| F06 | ST46 | 2021-01-09 | Panther Fusion | 27 | POSITIVE | 33.06 | 28.81 | POSITIVE
| C02 | ST11 | 2021-01-04 | Panther Fusion | 27.2 | POSITIVE | 36.05 | 28.92 | POSITIVE
| C07 | ST51 | 2021-01-07 | Panther Fusion | 28.5 | POSITIVE | 37.7 | 29.46 | POSITIVE
| G10 | ST79 | 2021-01-01 | Panther Fusion | 29 | POSITIVE | Und | 29.03 | NEGATIVE
| D05 | ST36 | 2021-01-05 | Panther Fusion | 29.6 | POSITIVE | 34.56 | 26.97 | POSITIVE
| A01 | ST01 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | Und | 30.97 | NEGATIVE
| B01 | ST02 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | Und | 27.42 | NEGATIVE
| C01 | ST03 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | Und | 30.04 | NEGATIVE
| E01 | ST05 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | 32.08 | NEGATIVE
| G01 | ST07 | 2021-01-14 | Panther Fusion | Und | NEGATIVE | Und | 30.23 | NEGATIVE
| H01 | ST08 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | 33.37 | NEGATIVE
| A02 | ST09 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | 27.92 | NEGATIVE
| E02 | ST13 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | 41.31 | 29.64 | NEGATIVE
| H02 | ST16 | 2021-01-14 | Panther Fusion | Und | NEGATIVE | Und | 34.26 | NEGATIVE
| A03 | ST17 | 2021-01-14 | Panther Fusion | Und | NEGATIVE | 42.58 | 28.19 | NEGATIVE
| C03 | ST19 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | 27.43 | NEGATIVE
| E03 | ST21 | 2021-01-14 | Panther Fusion | Und | NEGATIVE | Und | 33.75 | NEGATIVE
| B04 | ST26 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | 34.12 | NEGATIVE
| D04 | ST28 | 2021-01-03 | Panther Fusion | Und | NEGATIVE | Und | 28.02 | NEGATIVE
| G04 | ST31 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | Und | 33.19 | NEGATIVE
| H04 | ST32 | 2021-01-03 | Panther Fusion | Und | NEGATIVE | Und | 28.2 | NEGATIVE
| A05 | ST33 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | Und | 33.11 | NEGATIVE
| E05 | ST37 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | Und | 32.45 | NEGATIVE
| G05 | ST39 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | Und | 34.66 | NEGATIVE
| H05 | ST40 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | Und | 31.07 | NEGATIVE
| A06 | ST41 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | Und | 30.7 | NEGATIVE
| B06 | ST42 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | Und | 30.29 | NEGATIVE
| C06 | ST43 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | Und | 27.04 | NEGATIVE
| D06 | ST44 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | Und | 32.49 | NEGATIVE
| A07 | ST49 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | Und | 30.8 | NEGATIVE
| B07 | ST50 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | 26.1 | NEGATIVE
| F07 | ST54 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | 32.28 | NEGATIVE
| A08 | ST57 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | Und | 32.53 | NEGATIVE
| B08 | ST58 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | Und | 33.38 | NEGATIVE
| C08 | ST59 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | Und | 30.01 | NEGATIVE
| E08 | ST61 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | 33.07 | NEGATIVE
| F08 | ST62 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | Und | 29.73 | NEGATIVE
| C09 | ST67 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | 32.72 | NEGATIVE
| D09 | ST68 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | 29.25 | NEGATIVE
| F09 | ST70 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | 27.54 | NEGATIVE
| G09 | ST71 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | 27.88 | NEGATIVE
| A10 | ST73 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | 44.42 | 32.61 | NEGATIVE
| C10 | ST75 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | 32.92 | NEGATIVE
| E10 | ST77 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | 29.74 | NEGATIVE
| F10 | ST78 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | 31.53 | NEGATIVE
| B11 | Neg Control | N/A | N/A | N/A | N/A | Und | Und | NEGATIVE
| A11 | Pos Control | N/A | N/A | N/A | N/A | 34.15 | 32.1 | POSITIVE