## 1A LoD-Preliminary
1A LoD - Preliminary
Plate Map
|  | 1 | 2 | 3
| A Row | 100000 | 100000 | 100000
| B Row | 50000 | 50000 | 50000
| C Row | 25000 | 25000 | 25000
| D Row | 12500 | 12500 | 12500
| E Row | 6250 | 6250 | 6250
| F Row | 3125 | 3125 | 3125
| G Row | 0 | 0 | 0
| H Row | Neg Control | Neg Control | Pos Control
Results Image
| Line Item Data |  |  |  | Summary Data
| Concentration | Replicate | Result |  | Concentration | Result
| 100,00 copies/mL | 1 | POSITIVE |  | 100,00 copies/mL | POS 3/3
|  | 2 | POSITIVE |  | 50,000 copies/mL | POS 3/3
|  | 3 | POSITIVE |  | 25,000 copies/mL | POS 3/3
| 50,000 copies/mL | 1 | POSITIVE |  | 12,500 copies/mL | POS 3/3
|  | 2 | POSITIVE |  | 6,250 copies/mL | POS 3/3
|  | 3 | POSITIVE |  | 3,125 copies/mL | POS 3/3
| 25,000 copies/mL | 1 | POSITIVE |  | 0 copies/uL | NEG 3/3
|  | 2 | POSITIVE
|  | 3 | POSITIVE
| 12,500 copies/mL | 1 | POSITIVE
|  | 2 | POSITIVE
|  | 3 | POSITIVE
| 6,250 copies/mL | 1 | POSITIVE
|  | 2 | POSITIVE
|  | 3 | POSITIVE
| 3,125 copies/mL | 1 | POSITIVE
|  | 2 | POSITIVE
|  | 3 | POSITIVE
| 0 copies/mL | 1 | NEGATIVE
|  | 2 | NEGATIVE
|  | 3 | NEGATIVE
| Negative Control | 1 | NEGATIVE
|  | 2 | NEGATIVE
| Positive Control | 1 | POSITIVE

## 1B LoD-Confirmatory
1B LOD - Confirmatory
Supplemental Confirmatory LOD Run - *10% pipetting error on Colorimetric LAMP Master Mix
| Plate Map | [6 empty] | Plate Map
|  | 1 | 2 | 3 | 4 | 5 | 6 |  | 1 | 2 | 3 | 4 | 5 | 6
| A Row | 6250 | 6250 | 6250 | 12500 | 12500 | 12500 | A Row | 6250 | 6250 | 6250 | 12500 | 12500 | 12500
| B Row | 6250 | 6250 | 6250 | 12500 | 12500 | 12500 | B Row | 6250 | 6250 | 6250 | 12500 | 12500 | 12500
| C Row | 6250 | 6250 | 6250 | 12500 | 12500 | 12500 | C Row | 6250 | 6250 | 6250 | 12500 | 12500 | 12500
| D Row | 6250 | 6250 | 6250 | 12500 | 12500 | 12500 | D Row | 6250 | 6250 | 6250 | 12500 | 12500 | 12500
| E Row | 6250 | 6250 | 6250 | 12500 | 12500 | 12500 | E Row | 6250 | 6250 | 6250 | 12500 | 12500 | 12500
| F Row | 6250 | 6250 | 6250 | 12500 | 12500 | 12500 | F Row | 6250 | 6250 | 6250 | 12500 | 12500 | 12500
| G Row | 6250 | 6250 | 6250 | 12500 | 12500 | 12500 | G Row | 6250 | 6250 | 6250 | 12500 | 12500 | 12500
| H Row | 0 | Neg Control | Pos Control | 0 | Neg Control | Pos Control | H Row | 0 | Neg Control | Pos Control | 0 | Neg Control | Pos Control
| Results Image | [6 empty] | Results Image
| Line Item Data | 4: Summary Data | Line Item Data |  |  |  | Summary Data
| Concentration | Replicate | Result |  | Concentration | Result |  | Concentration | Replicate | Result |  | Concentration | Result
| 12,500 copies/mL | 1 | POSITIVE |  | 12,500 copies/mL | POS 20/21 |  | 12,500 copies/mL | 1 | POSITIVE |  | 12,500 copies/mL | POS 20/21
|  | 2 | POSITIVE |  | 6,250 copies/mL | POS 11/21 |  |  | 2 | POSITIVE |  | 6,250 copies/mL | POS 16/21
|  | 3 | POSITIVE |  | 0 copies/mL | NEG 2/2 |  |  | 3 | POSITIVE |  | 0 copies/mL | NEG 2/2
|  | 1: 4; 2: NEGATIVE |  | 4 | POSITIVE
|  | 1: 5; 2: POSITIVE |  | 5 | POSITIVE
|  | 1: 6; 2: POSITIVE |  | 6 | POSITIVE
|  | 1: 7; 2: POSITIVE |  | 7 | POSITIVE
|  | 1: 8; 2: POSITIVE |  | 8 | POSITIVE
|  | 1: 9; 2: POSITIVE |  | 9 | POSITIVE
|  | 1: 10; 2: POSITIVE |  | 10 | POSITIVE
|  | 1: 11; 2: POSITIVE |  | 11 | POSITIVE
|  | 1: 12; 2: POSITIVE |  | 12 | POSITIVE
|  | 1: 13; 2: POSITIVE |  | 13 | POSITIVE
|  | 1: 14; 2: POSITIVE |  | 14 | POSITIVE
|  | 1: 15; 2: POSITIVE |  | 15 | POSITIVE
|  | 1: 16; 2: POSITIVE |  | 16 | POSITIVE
|  | 1: 17; 2: POSITIVE |  | 17 | POSITIVE
|  | 1: 18; 2: POSITIVE |  | 18 | POSITIVE
|  | 1: 19; 2: POSITIVE |  | 19 | POSITIVE
|  | 1: 20; 2: POSITIVE |  | 20 | NEGATIVE
|  | 1: 21; 2: POSITIVE |  | 21 | POSITIVE
| 6,250 copies/mL | 1: 1; 2: POSITIVE | 6,250 copies/mL | 1 | POSITIVE
|  | 1: 2; 2: POSITIVE |  | 2 | POSITIVE
|  | 1: 3; 2: NEGATIVE |  | 3 | POSITIVE
|  | 1: 4; 2: POSITIVE |  | 4 | POSITIVE
|  | 1: 5; 2: POSITIVE |  | 5 | POSITIVE
|  | 1: 6; 2: NEGATIVE |  | 6 | NEGATIVE
|  | 1: 7; 2: POSITIVE |  | 7 | POSITIVE
|  | 1: 8; 2: NEGATIVE |  | 8 | POSITIVE
|  | 1: 9; 2: NEGATIVE |  | 9 | POSITIVE
|  | 1: 10; 2: NEGATIVE |  | 10 | POSITIVE
|  | 1: 11; 2: NEGATIVE |  | 11 | POSITIVE
|  | 1: 12; 2: NEGATIVE |  | 12 | POSITIVE
|  | 1: 13; 2: NEGATIVE |  | 13 | NEGATIVE
|  | 1: 14; 2: POSITIVE |  | 14 | POSITIVE
|  | 1: 15; 2: POSITIVE |  | 15 | POSITIVE
|  | 1: 16; 2: NEGATIVE |  | 16 | POSITIVE
|  | 1: 17; 2: POSITIVE |  | 17 | POSITIVE
|  | 1: 18; 2: POSITIVE |  | 18 | NEGATIVE
|  | 1: 19; 2: POSITIVE |  | 19 | POSITIVE
|  | 1: 20; 2: POSITIVE |  | 20 | NEGATIVE
|  | 1: 21; 2: NEGATIVE |  | 21 | NEGATIVE
| 0 copies/mL | 1: 1; 2: NEGATIVE | 0 copies/mL | 1 | NEGATIVE
|  | 1: 2; 2: NEGATIVE |  | 2 | NEGATIVE
| Negative Control | 1: 1; 2: NEGATIVE | Negative Control | 1 | NEGATIVE
|  | 1: 2; 2: NEGATIVE |  | 2 | NEGATIVE
| Positive Control | 1: 1; 2: POSITIVE | Positive Control | 1 | POSITIVE
|  | 1: 2; 2: POSITIVE |  | 2 | POSITIVE

## 2 Inclusivity - In Silico
2 Inclusivity (Analytical Sensitivity) - In Silico
All complete, high coverage SARS-CoV-2 sequences deposited at GISAID as of 2/27/2021
| Primer | AS1e (ORF1ab gene) |  | N2 (N gene) |  | E1 (E gene)
| Total Primer Length | 195 |  | 169 |  | 168
| Total Number of Strains Evaluated | 498224 |  | 498224 |  | 498224
| 100% Match | 474717 | 95.28% | 479548 | 96.25% | 462538 | 92.84%
| 1 Mismatch | 19301 | 3.874% | 15698 | 3.151% | 30626 | 6.147%
| 2 Mismatches | 338 | 0.06784% | 161 | 0.03231% | 1455 | 0.292%
| 3 Mismatches | 9 | 0.001806% | 5 | 0.001004% | 103 | 0.02067%
| > 3 Mismatches | 0 | 0% | 0 | 0% | 1 | 0.0002007%
| Total Strains Removed | 3859 | 0.7746% | 2812 | 0.5644% | 3501 | 0.7027%
| Removed_4+N | 2021 | 0.4056% | 1302 | 0.2613% | 1728 | 0.3468%
| Removed_2to4N | 357 | 0.07165% | 211 | 0.04235% | 291 | 0.05841%
| Removed_1N | 1481 | 0.2973% | 1299 | 0.2607% | 1482 | 0.2975%

## 3B Cross Reactivity - Wet Tests
3B Cross Reactivity (Analytical Specifity) - Wet Tests
| Line Item Data | [5 empty] | Plate Map
| Specimen ID | Replicate | Result |  |  |  |  | 1 | 2 | 3
| S82 | 1 | NEGATIVE |  |  |  | C Row | Pos Control | Pos Control | Neg Control
|  | 2 | NEGATIVE |  |  |  | D Row | S82 | S82 | S82
|  | 3 | NEGATIVE |  |  |  | E Row | M22 | M22 | M22
| M22 | 1 | NEGATIVE |  |  |  | F Row | R76 | R76 | R76
|  | 2 | NEGATIVE |  |  |  | G Row | I48 | I48 | I48
|  | 3 | NEGATIVE |  |  |  | H Row | H21 | H21 | H21
| R76 | 1 | NEGATIVE
|  | 2 | NEGATIVE |  |  |  | Results Image
|  | 3 | NEGATIVE
| I48 | 1 | NEGATIVE
|  | 2 | NEGATIVE
|  | 3 | NEGATIVE
| H21 | 1 | NEGATIVE
|  | 2 | NEGATIVE
|  | 3 | NEGATIVE
| Pos Control * | 1 | POSITIVE
|  | 2 | POSITIVE
| Negative Control | 1 | NEGATIVE
Sample Info
| Specimen ID | Item Number | Organism | Description | Strain | Batch No
| S82 | NR-3882 | SARS-CoV | UV-inactivated virus | N/A | 480405CA
| M22 | NR-49122 | Human Metapneumovirus | Genomic RNA from NR-22227 | TN/83-1211 | 359
| R76 | NR-43976 | RSV | Genomic RNA from NR-28527 | A1997/12-35 | 61727102
| I48 | NR-45848 | Influenza B | Genomic RNA from NR-44023 | Nevada/03/2011 | 70027288
| H21 | HM-121 | Streptococcus salivarius | Bacterial cell culture | SK126 | 70032423
Summary Data
|  |  | Organism | Description | Item Number | Detected Replicates
|  |  | SARS-CoV | UV-inactivated virus | NR-3882 | 0/3
|  |  | Human Metapneumovirus | Genomic RNA from NR-22227 | NR-49122 | 0/3
|  |  | RSV | Genomic RNA from NR-28527 | NR-43976 | 0/3
|  |  | Influenza B | Genomic RNA from NR-44023 | NR-45848 | 0/3
|  |  | Streptococcus salivarius | Bacterial cell culture | HM-121 | 0/3

## 3C Interfering Substances
3C Interfering Substances
| Plate Map | Spiked with irradiated virus |  |  | Unspiked
|  | 1 | 2 | 3 | 4 | 5 | 6
| A Row | INF1* | INF1* | INF1* | INF1 | INF1 | INF1
| B Row | INF2* | INF2* | INF2* | INF2 | INF2 | INF2
| C Row | INF3* | INF3* | INF3* | INF3 | INF3 | INF3
| D Row | INF4* | INF4* | INF4* | INF4 | INF4 | INF4
| E Row | INF5* | INF5* | INF5* | INF5 | INF5 | INF5
| F Row | INFNONE* | INFNONE* | INFNONE* | INFNONE | INFNONE | INFNONE
| G Row | Pos Control* | Pos Control* | Pos Control* | Pos Control* | Pos Control* | Pos Control*
| H Row | Neg Control | Neg Control | Neg Control | Neg Control | Neg Control | Neg Control
Results Image
| Specimen ID | Interferring Substance | Active Ingredient | Concentration | % Agreement with Expected Results
[5 empty] | Positive Control Spiked | Negative Control Unspiked
| INF1 | Blood | N/A | 1% v/v | 100% (3/3) | 100% (3/3)
| INF2 | Nasal Congestion Spray | Acetaminophen, Guaifenesin, Phenylephrine HCI | 20% v/v | 100% (3/3) | 100% (3/3)
| INF3 | Nasal Allergy Spray | Oxymetazoline HCl | 15% v/v | 100% (3/3) | 100% (3/3)
| INF4 | Lozenges | Menthol | 10% w/v | 100% (3/3) | 100% (3/3)
| INF5 | Mucin | N/A | 0.5% w/v | 100% (3/3) | 100% (3/3)
| Specimen ID | Interferring Substance | Spiked? | Replicate | Result
| INF1* | Blood | Spiked | 1 | POSITIVE
|  |  |  | 2 | POSITIVE
|  |  |  | 3 | POSITIVE
| INF1 | Blood | Unspiked | 1 | NEGATIVE
|  |  |  | 2 | NEGATIVE
|  |  |  | 3 | NEGATIVE
| INF2* | Nasal Congestion Spray | Spiked | 1 | POSITIVE
|  |  |  | 2 | POSITIVE
|  |  |  | 3 | POSITIVE
| INF2 | Nasal Congestion Spray | Unspiked | 1 | NEGATIVE
|  |  |  | 2 | NEGATIVE
|  |  |  | 3 | NEGATIVE
| INF3* | Nasal Allergy Spray | Spiked | 1 | POSITIVE
|  |  |  | 2 | POSITIVE
|  |  |  | 3 | POSITIVE
| INF3 | Nasal Allergy Spray | Unspiked | 1 | NEGATIVE
|  |  |  | 2 | NEGATIVE
|  |  |  | 3 | NEGATIVE
| INF4* | Lozenges | Spiked | 1 | POSITIVE
|  |  |  | 2 | POSITIVE
|  |  |  | 3 | POSITIVE
| INF4 | Lozenges | Unspiked | 1 | NEGATIVE
|  |  |  | 2 | NEGATIVE
|  |  |  | 3 | NEGATIVE
| INF5* | Mucin | Spiked | 1 | POSITIVE
|  |  |  | 2 | POSITIVE
|  |  |  | 3 | POSITIVE
| INF5 | Mucin | Unspiked | 1 | NEGATIVE
|  |  |  | 2 | NEGATIVE
|  |  |  | 3 | NEGATIVE
| INFNONE* | None | Spiked | 1 | POSITIVE
|  |  |  | 2 | POSITIVE
|  |  |  | 3 | POSITIVE
| INFNONE | None | Unspiked | 1 | NEGATIVE
|  |  |  | 2 | NEGATIVE
|  |  |  | 3 | NEGATIVE
| Positive Control* | None | Spiked | 1 | POSITIVE
|  |  |  | 2 | POSITIVE
|  |  |  | 3 | POSITIVE
|  |  |  | 4 | POSITIVE
|  |  |  | 5 | POSITIVE
|  |  |  | 6 | POSITIVE
| Negative Control | None | Unspiked | 1 | NEGATIVE
|  |  |  | 2 | NEGATIVE
|  |  |  | 3 | NEGATIVE
|  |  |  | 4 | NEGATIVE
|  |  |  | 5 | NEGATIVE
|  |  |  | 6 | NEGATIVE

## 4 Clinical Evaluation
4 Clinical Evaluation
Plate Map
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
| A Row | ST01 | ST09 | ST17 | ST25 | ST33 | ST41 | ST49 | ST57 | ST65 | ST73 | Pos Control
| B Row | ST02 | ST10 | ST18 | ST26 | ST34 | ST42 | ST50 | ST58 | ST66 | ST74 | Neg Control
| C Row | ST03 | ST11 | ST19 | ST27 | ST35 | ST43 | ST51 | ST59 | ST67 | ST75
| D Row | ST04 | ST12 | ST20 | ST28 | ST36 | ST44 | ST52 | ST60 | ST68 | ST76
| E Row | ST05 | ST13 | ST21 | ST29 | ST37 | ST45 | ST53 | ST61 | ST69 | ST77
| F Row | ST06 | ST14 | ST22 | ST30 | ST38 | ST46 | ST54 | ST62 | ST70 | ST78
| G Row | ST07 | ST15 | ST23 | ST31 | ST39 | ST47 | ST55 | ST63 | ST71 | ST79
| H Row | ST08 | ST16 | ST24 | ST32 | ST40 | ST48 | ST56 | ST64 | ST72 | ST80
Summary Data
| FloodLAMP QuickColor Test Result | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: Comparator Positive; 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: Comparator Negative; 3 ST17 ST18 ST19 ST20 ST21 ST22 ST23 ST24: Total
| Positive | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 36; 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: 0; 3 ST17 ST18 ST19 ST20 ST21 ST22 ST23 ST24: 36 |  |  | Results Image | Top: Saturation Filter Applied
| Negative | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 4; 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: 40; 3 ST17 ST18 ST19 ST20 ST21 ST22 ST23 ST24: 44 |  |  | Bottom: Original Image
| Total | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 40; 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: 40; 3 ST17 ST18 ST19 ST20 ST21 ST22 ST23 ST24: 80
| Positive Agreement | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 90.0% (36/40) 95% CI: 76.3% to 97.2%
| Negative Agreement | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 100% (40/40) 95% CI: 91.2% to 100%
Source of Specimens: Stanford COVID-19 Clinical Testing Program
Specimen Type: Anterior Nares Swab in PBS, previously tested and frozen
Comparator Test: Hologic Panther Fusion SARS-CoV-2 Assay and Hologic Panther Aptima SARS-CoV-2 Assay
* The FloodLAMP EasyPCR Test has been submitted as a companion EUA. Clinical Evaluation of both tests were performed together, using the same inactivated specimens.
Per the Instructions for Use, for the 7 inconclusive initial results, the follow up FloodLAMP EasyPCR COVID-19 Test was peformed with all results being positive
| Line Item Data | 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: Comparator - Hologic Panther; 6 ST41 ST42 ST43 ST44 ST45 ST46 ST47 ST48: FloodLAMP QuickColor Test; 8 ST57 ST58 ST59 ST60 ST61 ST62 ST63 ST64: FloodLAMP EasyPCR COVID-19 Test* | FloodLAMP QuickColor Test
| Well | Specimen ID | Comparator Test Date | Comparator Test Assay | Comparator Ct | Comparator Result | Initial Result |  | N1 Ct | RP Ct | Result | Final Interpretation
| D01 | ST04 | 2021-01-02 | Panther Fusion | 13.2 | POSITIVE | POSITIVE |  | 22.14 | Und | POSITIVE | POSITIVE
| C05 | ST35 | 2021-01-02 | Panther Fusion | 13.6 | POSITIVE | POSITIVE |  | 23.36 | Und | POSITIVE | POSITIVE
| H07 | ST56 | 2021-01-02 | Panther Fusion | 14 | POSITIVE | POSITIVE |  | 23.72 | Und | POSITIVE | POSITIVE
| G03 | ST23 | 2021-01-01 | Panther Fusion | 15 | POSITIVE | POSITIVE |  | 20.48 | Und | POSITIVE | POSITIVE
| C04 | ST27 | 2021-01-04 | Panther Fusion | 15.1 | POSITIVE | POSITIVE |  | 21.42 | Und | POSITIVE | POSITIVE
| E09 | ST69 | 2021-01-03 | Panther Fusion | 15.4 | POSITIVE | POSITIVE |  | 23.94 | 44.84 | POSITIVE | POSITIVE
| F02 | ST14 | 2021-01-03 | Panther Fusion | 15.5 | POSITIVE | POSITIVE |  | 22.92 | Und | POSITIVE | POSITIVE
| B02 | ST10 | 2021-01-01 | Panther Fusion | 15.6 | POSITIVE | POSITIVE |  | 27.68 | 29.88 | POSITIVE | POSITIVE
| E07 | ST53 | 2021-01-02 | Panther Fusion | 15.6 | POSITIVE | POSITIVE |  | 25.09 | 26.37 | POSITIVE | POSITIVE
| D08 | ST60 | 2021-01-03 | Panther Fusion | 15.7 | POSITIVE | POSITIVE |  | 21.65 | 39.12 | POSITIVE | POSITIVE
| A04 | ST25 | 2021-01-01 | Panther Fusion | 16.5 | POSITIVE | POSITIVE |  | 26.53 | 29.28 | POSITIVE | POSITIVE
| F03 | ST22 | 2021-01-03 | Panther Fusion | 17.2 | POSITIVE | POSITIVE |  | 24.73 | Und | POSITIVE | POSITIVE
| G02 | ST15 | 2021-01-01 | Panther Fusion | 17.4 | POSITIVE | POSITIVE |  | 29.13 | 28.91 | POSITIVE | POSITIVE
| B03 | ST18 | 2021-01-02 | Panther Fusion | 17.5 | POSITIVE | POSITIVE |  | 27.71 | 28.42 | POSITIVE | POSITIVE
| H03 | ST24 | 2021-01-01 | Panther Fusion | 18.9 | POSITIVE | POSITIVE |  | 30.06 | 29.81 | POSITIVE | POSITIVE
| E06 | ST45 | 2021-01-06 | Panther Fusion | 18.9 | POSITIVE | POSITIVE |  | 25.17 | 34.74 | POSITIVE | POSITIVE
| G08 | ST63 | 2021-01-02 | Panther Fusion | 19.4 | POSITIVE | POSITIVE |  | 26.19 | 27.69 | POSITIVE | POSITIVE
| D07 | ST52 | 2021-01-01 | Panther Fusion | 19.8 | POSITIVE | POSITIVE |  | 29.97 | 28.32 | POSITIVE | POSITIVE
| H06 | ST48 | 2021-01-02 | Panther Fusion | 20 | POSITIVE | POSITIVE |  | 32.66 | 32.43 | POSITIVE | POSITIVE
| E04 | ST29 | 2021-01-01 | Panther Fusion | 20.5 | POSITIVE | POSITIVE |  | 30.69 | 29.67 | POSITIVE | POSITIVE
| B09 | ST66 | 2021-01-01 | Panther Fusion | 21.1 | POSITIVE | POSITIVE |  | 28.81 | 28.38 | POSITIVE | POSITIVE
| F01 | ST06 | 2021-01-05 | Panther Fusion | 22.6 | POSITIVE | POSITIVE |  | 29.72 | 27.85 | POSITIVE | POSITIVE
| F05 | ST38 | 2021-01-01 | Panther Fusion | 23.2 | POSITIVE | POSITIVE |  | 32.03 | 27.71 | POSITIVE | POSITIVE
| F04 | ST30 | 2021-01-09 | Panther Fusion | 23.7 | POSITIVE | POSITIVE |  | 31.05 | 31.98 | POSITIVE | POSITIVE
| D10 | ST76 | 2021-01-02 | Panther Fusion | 24.1 | POSITIVE | POSITIVE |  | 32.4 | 28.55 | POSITIVE | POSITIVE
| G07 | ST55 | 2021-01-01 | Panther Fusion | 24.8 | POSITIVE | POSITIVE |  | 32.38 | 28.19 | POSITIVE | POSITIVE
| D03 | ST20 | 2021-01-01 | Panther Fusion | 25.1 | POSITIVE | INCONCLUSIVE |  | 35.05 | 26.4 | POSITIVE | POSITIVE
| H10 | ST80 | 2021-01-01 | Panther Fusion | 25.2 | POSITIVE | INCONCLUSIVE |  | 36.62 | 30.51 | POSITIVE | POSITIVE
| H08 | ST64 | 2021-01-02 | Panther Fusion | 25.6 | POSITIVE | INCONCLUSIVE |  | 37.94 | 32.77 | POSITIVE | POSITIVE
| H09 | ST72 | 2021-01-02 | Panther Fusion | 25.6 | POSITIVE | NEGATIVE |  | 37.26 | 31.28 | POSITIVE | NEGATIVE
| B10 | ST74 | 2021-01-02 | Panther Fusion | 25.7 | POSITIVE | INCONCLUSIVE |  | 35.56 | 31.43 | POSITIVE | POSITIVE
| G06 | ST47 | 2021-01-02 | Panther Fusion | 25.9 | POSITIVE | NEGATIVE |  | 36.48 | 30.08 | POSITIVE | NEGATIVE
| A09 | ST65 | 2021-01-05 | Panther Fusion | 26.1 | POSITIVE | POSITIVE |  | 33.86 | 30.38 | POSITIVE | POSITIVE
| D02 | ST12 | 2021-01-02 | Panther Fusion | 26.3 | POSITIVE | INCONCLUSIVE |  | 36.12 | 32.04 | POSITIVE | POSITIVE
| B05 | ST34 | 2021-01-04 | Panther Fusion | 26.4 | POSITIVE | INCONCLUSIVE |  | 33.78 | 28.91 | POSITIVE | POSITIVE
| F06 | ST46 | 2021-01-09 | Panther Fusion | 27 | POSITIVE | POSITIVE |  | 33.06 | 28.81 | POSITIVE | POSITIVE
| C02 | ST11 | 2021-01-04 | Panther Fusion | 27.2 | POSITIVE | INCONCLUSIVE |  | 36.05 | 28.92 | POSITIVE | POSITIVE
| C07 | ST51 | 2021-01-07 | Panther Fusion | 28.5 | POSITIVE | NEGATIVE |  | 37.7 | 29.46 | POSITIVE | NEGATIVE
| G10 | ST79 | 2021-01-01 | Panther Fusion | 29 | POSITIVE | NEGATIVE |  | Und | 29.03 | NEGATIVE | NEGATIVE
| D05 | ST36 | 2021-01-05 | Panther Fusion | 29.6 | POSITIVE | POSITIVE |  | 34.56 | 26.97 | POSITIVE | POSITIVE
| A01 | ST01 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | NEGATIVE |  | Und | 30.97 | NEGATIVE | NEGATIVE
| B01 | ST02 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | NEGATIVE |  | Und | 27.42 | NEGATIVE | NEGATIVE
| C01 | ST03 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | NEGATIVE |  | Und | 30.04 | NEGATIVE | NEGATIVE
| E01 | ST05 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 32.08 | NEGATIVE | NEGATIVE
| G01 | ST07 | 2021-01-14 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 30.23 | NEGATIVE | NEGATIVE
| H01 | ST08 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 33.37 | NEGATIVE | NEGATIVE
| A02 | ST09 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 27.92 | NEGATIVE | NEGATIVE
| E02 | ST13 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | 41.31 | 29.64 | NEGATIVE | NEGATIVE
| H02 | ST16 | 2021-01-14 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 34.26 | NEGATIVE | NEGATIVE
| A03 | ST17 | 2021-01-14 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | 42.58 | 28.19 | NEGATIVE | NEGATIVE
| C03 | ST19 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 27.43 | NEGATIVE | NEGATIVE
| E03 | ST21 | 2021-01-14 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 33.75 | NEGATIVE | NEGATIVE
| B04 | ST26 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 34.12 | NEGATIVE | NEGATIVE
| D04 | ST28 | 2021-01-03 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 28.02 | NEGATIVE | NEGATIVE
| G04 | ST31 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | NEGATIVE |  | Und | 33.19 | NEGATIVE | NEGATIVE
| H04 | ST32 | 2021-01-03 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 28.2 | NEGATIVE | NEGATIVE
| A05 | ST33 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 33.11 | NEGATIVE | NEGATIVE
| E05 | ST37 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 32.45 | NEGATIVE | NEGATIVE
| G05 | ST39 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 34.66 | NEGATIVE | NEGATIVE
| H05 | ST40 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 31.07 | NEGATIVE | NEGATIVE
| A06 | ST41 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 30.7 | NEGATIVE | NEGATIVE
| B06 | ST42 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 30.29 | NEGATIVE | NEGATIVE
| C06 | ST43 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | NEGATIVE |  | Und | 27.04 | NEGATIVE | NEGATIVE
| D06 | ST44 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | NEGATIVE |  | Und | 32.49 | NEGATIVE | NEGATIVE
| A07 | ST49 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | NEGATIVE |  | Und | 30.8 | NEGATIVE | NEGATIVE
| B07 | ST50 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 26.1 | NEGATIVE | NEGATIVE
| F07 | ST54 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 32.28 | NEGATIVE | NEGATIVE
| A08 | ST57 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | NEGATIVE |  | Und | 32.53 | NEGATIVE | NEGATIVE
| B08 | ST58 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | NEGATIVE |  | Und | 33.38 | NEGATIVE | NEGATIVE
| C08 | ST59 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | NEGATIVE |  | Und | 30.01 | NEGATIVE | NEGATIVE
| E08 | ST61 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 33.07 | NEGATIVE | NEGATIVE
| F08 | ST62 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | NEGATIVE |  | Und | 29.73 | NEGATIVE | NEGATIVE
| C09 | ST67 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 32.72 | NEGATIVE | NEGATIVE
| D09 | ST68 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 29.25 | NEGATIVE | NEGATIVE
| F09 | ST70 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 27.54 | NEGATIVE | NEGATIVE
| G09 | ST71 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 27.88 | NEGATIVE | NEGATIVE
| A10 | ST73 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | 44.42 | 32.61 | NEGATIVE | NEGATIVE
| C10 | ST75 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 32.92 | NEGATIVE | NEGATIVE
| E10 | ST77 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 29.74 | NEGATIVE | NEGATIVE
| F10 | ST78 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | NEGATIVE |  | Und | 31.53 | NEGATIVE | NEGATIVE
| B11 | Neg Control | N/A | N/A | N/A | N/A | NEGATIVE |  |  |  |  | NEGATIVE
| A11 | Pos Control | N/A | N/A | N/A | N/A | POSITIVE |  | 34.15 | 32.1 | POSITIVE | POSITIVE