## 1A LoD-Preliminary
1A LoD - Preliminary
Plate Map
|  | 1 | 2 | 3
| A Row | 100000 | 100000 | 100000
| B Row | 50000 | 50000 | 50000
| C Row | 25000 | 25000 | 25000
| D Row | 12500 | 12500 | 12500
| E Row | 6250 | 6250 | 6250
| F Row | 3125 | 3125 | 3125
| G Row | 0 | 0 | 0
| H Row |  | Neg Control | Pos Control
| Line Item Data - Applied Biosystems QuantStudio 7 Pro |  |  |  | Summary Data - Applied Biosystems QuantStudio 7 Pro
| Concentration | Replicate | Ct Value | Result | Concentration | Mean Ct | SD CT | Detected
| 100,00 copies/mL | 1 | 13.896 | POSITIVE | 100,00 copies/mL | 13.3926 | 1.0832 | 100% (3/3)
|  | 2 | 14.1325 | POSITIVE | 50,000 copies/mL | 12.9271 | 1.6626 | 100% (3/3)
|  | 3 | 12.1493 | POSITIVE | 25,000 copies/mL | 16.1979 | 3.8342 | 100% (3/3)
| 50,000 copies/mL | 1 | 14.7538 | POSITIVE | 12,500 copies/mL | 15.6999 | 1.7197 | 66% (2/3)
|  | 2 | 11.5023 | POSITIVE | 6,250 copies/mL | 16.4317 | n/a | 33% (1/3)
|  | 3 | 12.5252 | POSITIVE | 3,125 copies/mL | 13.4632 | n/a | 33% (1/3)
| 25,000 copies/mL | 1 | 12.7559 | POSITIVE
|  | 2 | 20.3304 | POSITIVE
|  | 3 | 15.5074 | POSITIVE
| 12,500 copies/mL | 1 | 14.4839 | POSITIVE
|  | 2 | Undetermined | NEGATIVE
|  | 3 | 16.9159 | POSITIVE
| 6,250 copies/mL | 1 | Undetermined | NEGATIVE
|  | 2 | Undetermined | NEGATIVE
|  | 3 | 16.4317 | POSITIVE
| 3,125 copies/mL | 1 | Undetermined | NEGATIVE
|  | 2 | Undetermined | NEGATIVE
|  | 3 | 13.4632 | POSITIVE
| 0 copies/mL | 1 | Undetermined | NEGATIVE
|  | 2 | Undetermined | NEGATIVE
|  | 3 | Undetermined | NEGATIVE
| Negative Control | 1 | Undetermined | NEGATIVE
| Positive Control | 1 | 11.5 | POSITIVE

## 1B LoD-Confirmatory
1A LoD - Confirmatory
Plate Map
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
| A Row | 50000 | 50000 | 50000 | 25000 | 25000 | 25000 | 12500 | 12500 | 12500
| B Row | 50000 | 50000 | 50000 | 25000 | 25000 | 25000 | 12500 | 12500 | 12500
| C Row | 50000 | 50000 | 50000 | 25000 | 25000 | 25000 | 12500 | 12500 | 12500
| D Row | 50000 | 50000 | 50000 | 25000 | 25000 | 25000 | 12500 | 12500 | 12500
| E Row | 50000 | 50000 | 50000 | 25000 | 25000 | 25000 | 12500 | 12500 | 12500
| F Row | 50000 | 50000 | 50000 | 25000 | 25000 | 25000 | 12500 | 12500 | 12500
| G Row | 50000 | 50000 | 50000 | 25000 | 25000 | 25000 | 12500 | 12500 | 12500
| H Row | 0 | 0 | 0 | Neg Control | Neg Control | Neg Control | Neg Control | Neg Control | Pos Control
| Line Item Data - Applied Biosystems QuantStudio 7 Pro | 5: Summary Data - Applied Biosystems QuantStudio 7 Pro
| Concentration | Replicate | Ct Value | Result |  | Concentration | Ct Mean | SD Ct | Detected
| 50,000 copies/mL | 1 | Undetermined | NEGATIVE |  | 50,000 copies/mL | 11.925 | 1.4821 | 95% (20/21)
|  | 2 | 10.3 | POSITIVE |  | 25,000 copies/mL | 13.8313 | 2.3392 | 76% (16/21)
|  | 3 | 14.9 | POSITIVE |  | 12,500 copies/mL | 14.725 | 3.2557 | 38% (8/21)
|  | 4 | 13.8 | POSITIVE
|  | 5 | 11.8 | POSITIVE
|  | 6 | 9.9 | POSITIVE
|  | 7 | 11.7 | POSITIVE
|  | 8 | 12.9 | POSITIVE
|  | 9 | 14.4 | POSITIVE
|  | 10 | 10.9 | POSITIVE
|  | 11 | 10.4 | POSITIVE
|  | 12 | 14.2 | POSITIVE
|  | 13 | 11.1 | POSITIVE
|  | 14 | 11.7 | POSITIVE
|  | 15 | 11.8 | POSITIVE
|  | 16 | 10.2 | POSITIVE
|  | 17 | 12.8 | POSITIVE
|  | 18 | 11 | POSITIVE
|  | 19 | 10.7 | POSITIVE
|  | 20 | 12.2 | POSITIVE
|  | 21 | 11.8 | POSITIVE
| 25,000 copies/mL | 1: 1; 2: 13; 3: POSITIVE
|  | 2 | 14.3 | POSITIVE
|  | 3 | 11.6 | POSITIVE
|  | 4 | Undetermined | NEGATIVE
|  | 5 | 13.7 | POSITIVE
|  | 6 | 13.1 | POSITIVE
|  | 7 | 11.3 | POSITIVE
|  | 8 | Undetermined | NEGATIVE
|  | 9 | 15.3 | POSITIVE
|  | 10 | 16.2 | POSITIVE
|  | 11 | 20.1 | POSITIVE
|  | 12 | 12.4 | POSITIVE
|  | 13 | 13.3 | POSITIVE
|  | 14 | 11.9 | POSITIVE
|  | 15 | 12.8 | POSITIVE
|  | 16 | 16.6 | POSITIVE
|  | 17 | Undetermined | NEGATIVE
|  | 18 | 14.5 | POSITIVE
|  | 19 | Undetermined | NEGATIVE
|  | 20 | 11.2 | POSITIVE
|  | 21 | Undetermined | NEGATIVE
| 12,500 copies/mL | 1: 1; 2: 19; 3: POSITIVE
|  | 2 | Undetermined | NEGATIVE
|  | 3 | Undetermined | NEGATIVE
|  | 4 | 15.7 | POSITIVE
|  | 5 | Undetermined | NEGATIVE
|  | 6 | 11.8 | POSITIVE
|  | 7 | 11 | POSITIVE
|  | 8 | Undetermined | NEGATIVE
|  | 9 | 14.1 | POSITIVE
|  | 10 | Undetermined | NEGATIVE
|  | 11 | Undetermined | NEGATIVE
|  | 12 | Undetermined | NEGATIVE
|  | 13 | Undetermined | NEGATIVE
|  | 14 | Undetermined | NEGATIVE
|  | 15 | 18.1 | POSITIVE
|  | 16 | Undetermined | NEGATIVE
|  | 17 | Undetermined | NEGATIVE
|  | 18 | Undetermined | NEGATIVE
|  | 19 | Undetermined | NEGATIVE
|  | 20 | 17.2 | POSITIVE
|  | 21 | 10.9 | POSITIVE
| 0 copies/mL | 1: 1; 2: Undetermined; 3: NEGATIVE
|  | 2 | Undetermined | NEGATIVE
|  | 3 | Undetermined | NEGATIVE
| Negative Control | 1: 1; 2: Undetermined; 3: NEGATIVE
|  | 2 | Undetermined | NEGATIVE
|  | 3 | Undetermined | NEGATIVE
|  | 4 | Undetermined | NEGATIVE
|  | 5 | Undetermined | NEGATIVE
| Positive Control | 1: 1; 2: 10.5; 3: POSITIVE

## Inclusivity - In Silico
2 Inclusivity (Analytical Sensitivity) - In Silico
All complete, high coverage SARS-CoV-2 sequences deposited at GISAID as of 2/27/2021
| Primer | AS1e (ORF1ab gene) |  | N2 (N gene) |  | E1 (E gene)
| Total Primer Length | 195 |  | 169 |  | 168
| Total Number of Strains Evaluated | 498224 |  | 498224 |  | 498224
| 100% Match | 474717 | 95.28% | 479548 | 96.25% | 462538 | 92.84%
| 1 Mismatch | 19301 | 3.874% | 15698 | 3.151% | 30626 | 6.147%
| 2 Mismatches | 338 | 0.06784% | 161 | 0.03231% | 1455 | 0.292%
| 3 Mismatches | 9 | 0.001806% | 5 | 0.001004% | 103 | 0.02067%
| > 3 Mismatches | 0 | 0% | 0 | 0% | 1 | 0.0002007%
| Total Strains Removed | 3859 | 0.7746% | 2812 | 0.5644% | 3501 | 0.7027%
| Removed_4+N | 2021 | 0.4056% | 1302 | 0.2613% | 1728 | 0.3468%
| Removed_2to4N | 357 | 0.07165% | 211 | 0.04235% | 291 | 0.05841%
| Removed_1N | 1481 | 0.2973% | 1299 | 0.2607% | 1482 | 0.2975%

## 3B Cross Reactivity - Wet Tests
3B Cross Reactivity (Analytical Specifity) - Wet Tests
| Line Item Data - Applied Biosystems QuantStudio 7 Pro |  |  |  |  | Plate Map
| Specimen ID | Replicate | Ct Value | Result |  |  | 4 | 5 | 6 | 7
| S82 | 1 | Undetermined | NEGATIVE |  | A Row |  |  |  | H21
|  | 2 | Undetermined | NEGATIVE |  | B Row |  |  |  | H21
|  | 3 | Undetermined | NEGATIVE |  | C Row |  |  |  | H21
| M22 | 1 | Undetermined | NEGATIVE |  | D Row |  |  |  | Neg Control
|  | 2 | Undetermined | NEGATIVE |  | E Row | S82 | S82 | S82 | Neg Control
|  | 3 | Undetermined | NEGATIVE |  | F Row | M22 | M22 | M22 | Pos Control
| R76 | 1 | Undetermined | NEGATIVE |  | G Row | R76 | R76 | R76 | Pos Control
|  | 2 | Undetermined | NEGATIVE |  | H Row | I48 | I48 | I48 | Pos Control
|  | 3 | Undetermined | NEGATIVE
| I48 | 1 | Undetermined | NEGATIVE
|  | 2 | Undetermined | NEGATIVE
|  | 3 | Undetermined | NEGATIVE
| H21 | 1 | Undetermined | NEGATIVE
|  | 2 | Undetermined | NEGATIVE
|  | 3 | Undetermined | NEGATIVE
| Pos Control | 1 | 14.4716 | POSITIVE
|  | 2 | 15.2636 | POSITIVE
|  | 3 | 15.0132 | POSITIVE
| Negative Control | 1 | Undetermined | NEGATIVE
|  | 2 | Undetermined | NEGATIVE
Sample Info
| Specimen ID | Item Number | Organism | Description | Strain | Batch No
| S82 | NR-3882 | SARS-CoV | UV-inactivated virus | N/A | 480405CA
| M22 | NR-49122 | Human Metapneumovirus | Genomic RNA from NR-22227 | TN/83-1211 | 359
| R76 | NR-43976 | RSV | Genomic RNA from NR-28527 | A1997/12-35 | 61727102
| I48 | NR-45848 | Influenza B | Genomic RNA from NR-44023 | Nevada/03/2011 | 70027288
| H21 | HM-121 | Streptococcus salivarius | Bacterial cell culture | SK126 | 70032423
Summary Data - Applied Biosystems QuantStudio 7 Pro
|  |  | Organism | Description | Item Number | Detected Replicates
|  |  | SARS-CoV | UV-inactivated virus | NR-3882 | 0/3
|  |  | Human Metapneumovirus | Genomic RNA from NR-22227 | NR-49122 | 0/3
|  |  | RSV | Genomic RNA from NR-28527 | NR-43976 | 0/3
|  |  | Influenza B | Genomic RNA from NR-44023 | NR-45848 | 0/3
|  |  | Streptococcus salivarius | Bacterial cell culture | HM-121 | 0/3

## 3C Interfering Substances
3C Interfering Substances
Plate Map
|  | 1 | 2 | 3 | 4 | 5 | 6
| A Row | INF1* | INF1* | INF1* | INF1 | INF1 | INF1
| B Row | INF3* | INF3* | INF3* | INF3 | INF3 | INF3
| C Row | INF4* | INF4* | INF4* | INF4 | INF4 | INF4
| D Row | INF9* | INF9* | INF9* | INF9 | INF9 | INF9
| E Row | INF10* | INF10* | INF10* | INF10 | INF10 | INF10
| F Row | NOINF* | NOINF* | NOINF* |  |  | Pos Control*
| G Row | 1: Neg Control; 2: Neg Control; 3: Neg Control
Summary Data - Applied Biosystems QuantStudio 7 Pro
| Specimen ID | Interferring Substance | Active Ingredient | Concentration | Positive Control Spiked |  | Negative Control Unspiked |  | % Agreement with Expected Results
|  | 4: Mean Ct; 5: SD Ct; 6: Mean Ct | SD Ct | Positive Control Spiked | Negative Control Unspiked
| INF1 | Blood | N/A | 1% v/v | 12.6218 | 2.1725 | Undetermined | n/a | 100% (3/3) | 100% (3/3)
| INF2 | Nasal Congestion Spray | Acetaminophen, Guaifenesin, Phenylephrine HCI | 20% v/v | 11.2233 | 0.5846 | Undetermined | n/a | 100% (3/3) | 100% (3/3)
| INF3 | Nasal Allergy Spray | Oxymetazoline HCl | 15% v/v | 11.3766 | 0.070037 | Undetermined | n/a | 100% (3/3) | 100% (3/3)
| INF4 | Lozenges | Menthol | 10% w/v | 13.5396 | 0.28212 | Undetermined | n/a | 100% (3/3) | 100% (3/3)
| INF5 | Mucin | N/A | 0.5% w/v | 11.5916 | 2.1647 | Undetermined | n/a | 100% (3/3) | 100% (3/3)
Line Item Data - Applied Biosystems QuantStudio 7 Pro
| Specimen ID | Interferring Substance | Spiked? | Replicate | Ct Value | Result
| INF1* | Blood | Spiked | 1 | 10.7133 | POSITIVE
|  |  |  | 2 | 14.986 | POSITIVE
|  |  |  | 3 | 12.1662 | POSITIVE
| INF1 | Blood | Unspiked | 1 | Undetermined | NEGATIVE
|  |  |  | 2 | Undetermined | NEGATIVE
|  |  |  | 3 | Undetermined | NEGATIVE
| INF2* | Nasal Congestion Spray | Spiked | 1 | 11.8144 | POSITIVE
|  |  |  | 2 | 11.2101 | POSITIVE
|  |  |  | 3 | 10.6454 | POSITIVE
| INF2 | Nasal Congestion Spray | Unspiked | 1 | Undetermined | NEGATIVE
|  |  |  | 2 | Undetermined | NEGATIVE
|  |  |  | 3 | Undetermined | NEGATIVE
| INF3* | Nasal Allergy Spray | Spiked | 1 | 11.3026 | POSITIVE
|  |  |  | 2 | 11.3853 | POSITIVE
|  |  |  | 3 | 11.4418 | POSITIVE
| INF3 | Nasal Allergy Spray | Unspiked | 1 | Undetermined | NEGATIVE
|  |  |  | 2 | Undetermined | NEGATIVE
|  |  |  | 3 | Undetermined | NEGATIVE
| INF4* | Lozenges | Spiked | 1 | 13.8651 | POSITIVE
|  |  |  | 2 | 13.3657 | POSITIVE
|  |  |  | 3 | 13.3879 | POSITIVE
| INF4 | Lozenges | Unspiked | 1 | Undetermined | NEGATIVE
|  |  |  | 2 | Undetermined | NEGATIVE
|  |  |  | 3 | Undetermined | NEGATIVE
| INF5* | Mucin | Spiked | 1 | 10.85 | POSITIVE
|  |  |  | 2 | 9.8952 | POSITIVE
|  |  |  | 3 | 14.0296 | POSITIVE
| INF5 | Mucin | Unspiked | 1 | Undetermined | NEGATIVE
|  |  |  | 2 | Undetermined | NEGATIVE
|  |  |  | 3 | Undetermined | NEGATIVE
| INFNONE* | None | Spiked | 1 | 10.85 | POSITIVE
|  |  |  | 2 | 9.8952 | POSITIVE
|  |  |  | 3 | 14.0296 | POSITIVE
| INFNONE | None | Unspiked | 1 | Undetermined | NEGATIVE
|  |  |  | 2 | Undetermined | NEGATIVE
|  |  |  | 3 | Undetermined | NEGATIVE
| Positive Control* | None | Spiked | 1 | 10.5 | POSITIVE
| Negative Control | None | Unspiked | 1 | Undetermined | NEGATIVE
|  |  |  | 2 | Undetermined | NEGATIVE
|  |  |  | 3 | Undetermined | NEGATIVE
|  |  |  | 4 | Undetermined | NEGATIVE

## 4 Clinical Evaluation
4 Clinical Evaluation
Plate Map
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
| A Row | ST01 | ST09 | ST17 | ST25 | ST33 | ST41 | ST49 | ST57 | ST65 | ST73 | Pos Control
| B Row | ST02 | ST10 | ST18 | ST26 | ST34 | ST42 | ST50 | ST58 | ST66 | ST74 | Neg Control
| C Row | ST03 | ST11 | ST19 | ST27 | ST35 | ST43 | ST51 | ST59 | ST67 | ST75
| D Row | ST04 | ST12 | ST20 | ST28 | ST36 | ST44 | ST52 | ST60 | ST68 | ST76
| E Row | ST05 | ST13 | ST21 | ST29 | ST37 | ST45 | ST53 | ST61 | ST69 | ST77
| F Row | ST06 | ST14 | ST22 | ST30 | ST38 | ST46 | ST54 | ST62 | ST70 | ST78
| G Row | ST07 | ST15 | ST23 | ST31 | ST39 | ST47 | ST55 | ST63 | ST71 | ST79
| H Row | ST08 | ST16 | ST24 | ST32 | ST40 | ST48 | ST56 | ST64 | ST72 | ST80
Summary Data
| FloodLAMP QuickFluor Result | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: Comparator Positive; 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: Comparator Negative; 3 ST17 ST18 ST19 ST20 ST21 ST22 ST23 ST24: Total
| Positive | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 34; 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: 0; 3 ST17 ST18 ST19 ST20 ST21 ST22 ST23 ST24: 34
| Negative | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 6; 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: 40; 3 ST17 ST18 ST19 ST20 ST21 ST22 ST23 ST24: 46
| Total | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 40; 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: 40; 3 ST17 ST18 ST19 ST20 ST21 ST22 ST23 ST24: 80
| Positive Agreement | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 85.0% (34/40) 95% CI: 70.2% to 94.3%
| Negative Agreement | 1 ST01 ST02 ST03 ST04 ST05 ST06 ST07 ST08: 100% (40/40) 95% CI: 91.2% to 100%
Source of Specimens: Stanford COVID-19 Clinical Testing Program
Specimen Type: Anterior Nares Swab in PBS, previously tested and frozen
Comparator Test: Hologic Panther Fusion SARS-CoV-2 Assay and Hologic Panther Aptima SARS-CoV-2 Assay
| Line Item Data | 2 ST09 ST10 ST11 ST12 ST13 ST14 ST15 ST16: Comparator - Hologic Panther; 6 ST41 ST42 ST43 ST44 ST45 ST46 ST47 ST48: FloodLAMP QuickFluor Test - Bio-Rad CFX96
| Well | Specimen ID | Comparator Test Date | Comparator Test Assay | Comparator Ct | Comparator Result | SARS Ct | Result
| D01 | ST04 | 2021-01-02 | Panther Fusion | 13.2 | POSITIVE | 8.09 | POSITIVE
| C05 | ST35 | 2021-01-02 | Panther Fusion | 13.6 | POSITIVE | 8.13 | POSITIVE
| H07 | ST56 | 2021-01-02 | Panther Fusion | 14 | POSITIVE | 8.29 | POSITIVE
| G03 | ST23 | 2021-01-01 | Panther Fusion | 15 | POSITIVE | 8.3 | POSITIVE
| C04 | ST27 | 2021-01-04 | Panther Fusion | 15.1 | POSITIVE | 9.25 | POSITIVE
| E09 | ST69 | 2021-01-03 | Panther Fusion | 15.4 | POSITIVE | 8.29 | POSITIVE
| F02 | ST14 | 2021-01-03 | Panther Fusion | 15.5 | POSITIVE | 8.12 | POSITIVE
| B02 | ST10 | 2021-01-01 | Panther Fusion | 15.6 | POSITIVE | 11.34 | POSITIVE
| E07 | ST53 | 2021-01-02 | Panther Fusion | 15.6 | POSITIVE | 8.64 | POSITIVE
| D08 | ST60 | 2021-01-03 | Panther Fusion | 15.7 | POSITIVE | 9.08 | POSITIVE
| A04 | ST25 | 2021-01-01 | Panther Fusion | 16.5 | POSITIVE | 10.93 | POSITIVE
| F03 | ST22 | 2021-01-03 | Panther Fusion | 17.2 | POSITIVE | 8.57 | POSITIVE
| G02 | ST15 | 2021-01-01 | Panther Fusion | 17.4 | POSITIVE | 10.09 | POSITIVE
| B03 | ST18 | 2021-01-02 | Panther Fusion | 17.5 | POSITIVE | 10.33 | POSITIVE
| H03 | ST24 | 2021-01-01 | Panther Fusion | 18.9 | POSITIVE | 10.47 | POSITIVE
| E06 | ST45 | 2021-01-06 | Panther Fusion | 18.9 | POSITIVE | 8.86 | POSITIVE
| G08 | ST63 | 2021-01-02 | Panther Fusion | 19.4 | POSITIVE | 11.18 | POSITIVE
| D07 | ST52 | 2021-01-01 | Panther Fusion | 19.8 | POSITIVE | 12.28 | POSITIVE
| H06 | ST48 | 2021-01-02 | Panther Fusion | 20 | POSITIVE | 12.67 | POSITIVE
| E04 | ST29 | 2021-01-01 | Panther Fusion | 20.5 | POSITIVE | 13.04 | POSITIVE
| B09 | ST66 | 2021-01-01 | Panther Fusion | 21.1 | POSITIVE | 11.5 | POSITIVE
| F01 | ST06 | 2021-01-05 | Panther Fusion | 22.6 | POSITIVE | 13.96 | POSITIVE
| F05 | ST38 | 2021-01-01 | Panther Fusion | 23.2 | POSITIVE | 11.1 | POSITIVE
| F04 | ST30 | 2021-01-09 | Panther Fusion | 23.7 | POSITIVE | 13.03 | POSITIVE
| D10 | ST76 | 2021-01-02 | Panther Fusion | 24.1 | POSITIVE | 14.26 | POSITIVE
| G07 | ST55 | 2021-01-01 | Panther Fusion | 24.8 | POSITIVE | 13.42 | POSITIVE
| D03 | ST20 | 2021-01-01 | Panther Fusion | 25.1 | POSITIVE | Und | NEGATIVE
| H10 | ST80 | 2021-01-01 | Panther Fusion | 25.2 | POSITIVE | 17.74 | POSITIVE
| H08 | ST64 | 2021-01-02 | Panther Fusion | 25.6 | POSITIVE | 14.31 | POSITIVE
| H09 | ST72 | 2021-01-02 | Panther Fusion | 25.6 | POSITIVE | Und | NEGATIVE
| B10 | ST74 | 2021-01-02 | Panther Fusion | 25.7 | POSITIVE | 14.24 | POSITIVE
| G06 | ST47 | 2021-01-02 | Panther Fusion | 25.9 | POSITIVE | Und | NEGATIVE
| A09 | ST65 | 2021-01-05 | Panther Fusion | 26.1 | POSITIVE | 13.05 | POSITIVE
| D02 | ST12 | 2021-01-02 | Panther Fusion | 26.3 | POSITIVE | Und | NEGATIVE
| B05 | ST34 | 2021-01-04 | Panther Fusion | 26.4 | POSITIVE | 14.32 | POSITIVE
| F06 | ST46 | 2021-01-09 | Panther Fusion | 27 | POSITIVE | 14.79 | POSITIVE
| C02 | ST11 | 2021-01-04 | Panther Fusion | 27.2 | POSITIVE | 15.8 | POSITIVE
| C07 | ST51 | 2021-01-07 | Panther Fusion | 28.5 | POSITIVE | Und | NEGATIVE
| G10 | ST79 | 2021-01-01 | Panther Fusion | 29 | POSITIVE | Und | NEGATIVE
| D05 | ST36 | 2021-01-05 | Panther Fusion | 29.6 | POSITIVE | 12.89 | POSITIVE
| A01 | ST01 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | Und | NEGATIVE
| B01 | ST02 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | Und | NEGATIVE
| C01 | ST03 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | Und | NEGATIVE
| E01 | ST05 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| G01 | ST07 | 2021-01-14 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| H01 | ST08 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| A02 | ST09 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| E02 | ST13 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| H02 | ST16 | 2021-01-14 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| A03 | ST17 | 2021-01-14 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| C03 | ST19 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| E03 | ST21 | 2021-01-14 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| B04 | ST26 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| D04 | ST28 | 2021-01-03 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| G04 | ST31 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | Und | NEGATIVE
| H04 | ST32 | 2021-01-03 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| A05 | ST33 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| E05 | ST37 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| G05 | ST39 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| H05 | ST40 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| A06 | ST41 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| B06 | ST42 | 2021-01-05 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| C06 | ST43 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | Und | NEGATIVE
| D06 | ST44 | 2021-01-05 | Panther Aptima | Und | NEGATIVE | Und | NEGATIVE
| A07 | ST49 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | Und | NEGATIVE
| B07 | ST50 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| F07 | ST54 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| A08 | ST57 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | Und | NEGATIVE
| B08 | ST58 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | Und | NEGATIVE
| C08 | ST59 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | Und | NEGATIVE
| E08 | ST61 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| F08 | ST62 | 2021-01-04 | Panther Aptima | Und | NEGATIVE | Und | NEGATIVE
| C09 | ST67 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| D09 | ST68 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| F09 | ST70 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| G09 | ST71 | 2021-01-04 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| A10 | ST73 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| C10 | ST75 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| E10 | ST77 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| F10 | ST78 | 2021-01-13 | Panther Fusion | Und | NEGATIVE | Und | NEGATIVE
| B11 | Neg Control | N/A | N/A | N/A | N/A | Und | NEGATIVE
| A11 | Pos Control | N/A | N/A | N/A | N/A | 10.88 | POSITIVE